## Sheet1
DATA released for validation of HyperOCR (NOAA); this is not a calibration of this instrument
OL455 Radiances September 19, 2012
Level 5 == micrometer at 5.0
Average of two methods of interpolating in NPR-based calibration using the VXR
I. polynomial for spectral scaling from NPR - Carol and Steph - file ol455_lamps.txt
II. Blackbody fit by Mike Feinholz - file OL455_bbfits_ol.txt
VXR has six channels from 410 nm to 870 nm
results reported in this range only
reference documents: appendixB.docx (NIST's FY12 report to NOAA); Calibration of NIST OL455_v2_with fits.docx
Wavelength | Spectral Radiance | Wavelength | Uncertainty
nm | uW/cm2/sr/nm | nm | %, k=2
410 | 0.96505 | 411.6427 | 1.9343
420 | 1.1922 | 440.9736 | 1.6503
430 | 1.4518 | 548.2224 | 1.4545
440 | 1.7432 | 661.0952 | 1.3274
450 | 2.0651 | 775.1697 | 1.2729
460 | 2.4222 | 869.5366 | 1.2535
470 | 2.81
480 | 3.226
490 | 3.6738
500 | 4.147
510 | 4.6445
520 | 5.1633
530 | 5.7067
540 | 6.2737
550 | 6.8354
560 | 7.4152
570 | 8.0075
580 | 8.6008
590 | 9.1944
600 | 9.7973
610 | 10.3935
620 | 10.986
630 | 11.5744
640 | 12.1478
650 | 12.7213
660 | 13.279
670 | 13.824
680 | 14.3722
690 | 14.8917
700 | 15.4013
710 | 15.9048
720 | 16.3723
730 | 16.8262
740 | 17.2739
750 | 17.6932
760 | 18.0539
770 | 18.4763
780 | 18.8327
790 | 19.1757
800 | 19.5357
810 | 19.8595
820 | 20.1468
830 | 20.4356
840 | 20.7026
850 | 20.9474
860 | 21.1928
870 | 21.423