## Sheet1
| Corrected Spectrum | [10 empty] | Radiance Calibration Value=(Tested Light Source Signal*Tfilter 1/Nipper Jr Signal)*(FASCAL Calibrated Nipper Radiance Value)
| CAS091228A | CAS091228B | CAS091228C | CAS091228D | CAS091230A | CAS091230B | CAS091230C | CAS091230D |  |  |  | CAS091228A | CAS091228B | CAS091228C | CAS091228D | CAS091230A | CAS091230B | CAS091230C
| OL420_No filter | OL420_BG39 | OL420_BG28 | OL420_PER100 | White LED/Channel 1_200 | Blue LED/Channel 2_200 | HeNe laser | Nipper Jr. with filter 1 |  |  |  | OL420_No filter | OL420_BG39 | OL420_BG28 | OL420_PER100 | White LED/Channel 1_200 | Blue LED/Channel 2_200 | HeNe laser
Integration Time (ms) | 95 | 200 | 650 | 155 | 80 | 18 | 45 | 9 |  | Cubic Spline fitted radiances (Radiance of Nipper Jr.) | Integration Time (ms) | 95 | 200 | 650 | 155 | 80 | 18 | 45
Wavelength ([nm] | Corrected Spectrum | Corrected Spectrum | Corrected Spectrum | Corrected Spectrum | Corrected Spectrum | Corrected Spectrum | Corrected Spectrum | Corrected Spectrum | Transmission of filter 1 |  | Wavelength ([nm] | Radiance(uW cm-2 sr-1 nm-1) | Radiance(uW cm-2 sr-1 nm-1) | Radiance(uW cm-2 sr-1 nm-1) | Radiance(uW cm-2 sr-1 nm-1) | Radiance(uW cm-2 sr-1 nm-1) | Radiance(uW cm-2 sr-1 nm-1) | Radiance(uW cm-2 sr-1 nm-1)
379.7622 | 0.047465 | 0.003012 | -0.001364 | -0.001501 | 0.003881 | 0.0421 | -0.008263 | 0.57 | 0.097009 | 4.8686 | 379.7622 | 4.1792 | 0.2652 | -0.1201 | -0.13216 | 0.34171 | 3.7068 | -0.72754
380.4498 | 0.047617 | 0.000674 | -0.000799 | -0.000257 | 0.005629 | 0.024878 | 0.002596 | 0.57222 | 0.049154 | 4.9512 | 380.4498 | 8.382 | 0.11864 | -0.14065 | -0.04524 | 0.99087 | 4.3793 | 0.45697
381.1374 | 0.045691 | 0.001829 | 0.000301 | -0.002905 | 0.001472 | 0.022735 | 0.001005 | 0.55667 | 0.10173 | 5.035 | 381.1374 | 4.0625 | 0.16262 | 0.026763 | -0.25829 | 0.13088 | 2.0214 | 0.089357
381.8249 | 0.055093 | 0.001024 | 1.6e-05 | -0.000761 | 0.003649 | 0.020339 | -0.000589 | 0.57956 | 0.034169 | 5.1202 | 381.8249 | 14.2445 | 0.26476 | 0.0041368 | -0.19676 | 0.94346 | 5.2587 | -0.15229
382.5125 | 0.049429 | 0.001265 | 0.001762 | 0.001175 | 0.006033 | 0.035762 | -0.005811 | 0.58888 | 0.074732 | 5.2066 | 382.5125 | 5.8479 | 0.14966 | 0.20846 | 0.13901 | 0.71376 | 4.231 | -0.6875
383.2001 | 0.051752 | 0.000813 | 0.001378 | -0.001121 | 0.003777 | 0.030458 | -0.002659 | 0.60279 | 0.039475 | 5.2942 | 383.2001 | 11.5144 | 0.18089 | 0.30659 | -0.24941 | 0.84035 | 6.7767 | -0.59161
383.8877 | 0.051265 | 0.002679 | 0.000915 | 0.00052 | 0.002632 | 0.026495 | 0.00228 | 0.59291 | 0.023015 | 5.383 | 383.8877 | 20.2231 | 1.0568 | 0.36095 | 0.20513 | 1.0383 | 10.4518 | 0.89942
384.5753 | 0.055585 | 0.001671 | 0.00079 | 0.003076 | 0.004518 | 0.021601 | 0.002877 | 0.62781 | 0.076399 | 5.473 | 384.5753 | 6.3426 | 0.19067 | 0.090143 | 0.35099 | 0.51553 | 2.4648 | 0.32828
385.2628 | 0.053095 | 0.00167 | 0.000416 | 0.001082 | 0.00613 | 0.020945 | -0.000439 | 0.63495 | 0.017525 | 5.5641 | 385.2628 | 26.5482 | 0.83502 | 0.20801 | 0.54101 | 3.0651 | 10.4728 | -0.21951
385.9504 | 0.054669 | 0.001754 | 0.002205 | -0.000659 | 0.004023 | 0.020063 | 0.000699 | 0.61437 | 0.028867 | 5.6563 | 385.9504 | 17.4354 | 0.5594 | 0.70323 | -0.21017 | 1.283 | 6.3986 | 0.22293
386.638 | 0.052478 | 0.002945 | -1.5e-05 | 0.00028 | 0.006081 | 0.005328 | -0.007567 | 0.62718 | 0.11181 | 5.7495 | 386.638 | 4.3028 | 0.24147 | -0.0012299 | 0.022958 | 0.49859 | 0.43685 | -0.62043
387.3256 | 0.056499 | 0.002744 | 0.001477 | -0.000641 | 0.008118 | 0.011907 | 0.001098 | 0.62443 | 0.033537 | 5.8438 | 387.3256 | 15.7664 | 0.76573 | 0.41216 | -0.17887 | 2.2654 | 3.3227 | 0.3064
388.0131 | 0.047453 | 0.002065 | 0.001682 | 0.000401 | 0.007194 | 0.017717 | -0.006527 | 0.66624 | 0.085512 | 5.9392 | 388.0131 | 4.9469 | 0.21527 | 0.17535 | 0.041803 | 0.74996 | 1.847 | -0.68043
388.7007 | 0.053664 | 0.000301 | 0.001128 | 0.001493 | 0.007406 | 0.034796 | 0.004607 | 0.61109 | 0.0071605 | 6.0355 | 388.7007 | 74.0193 | 0.41517 | 1.5559 | 2.0593 | 10.2152 | 47.9945 | 6.3545
389.3883 | 0.054885 | 0.002862 | 0.001528 | -0.001445 | 0.006593 | 0.041758 | -0.003705 | 0.6376 | 0.11719 | 6.1327 | 389.3883 | 4.5047 | 0.2349 | 0.12541 | -0.1186 | 0.54112 | 3.4273 | -0.30409
390.0759 | 0.052677 | 0.002812 | 0.000717 | -0.000978 | 0.003112 | 0.020036 | -0.001314 | 0.66309 | 0.071207 | 6.2309 | 390.0759 | 6.9515 | 0.37108 | 0.094618 | -0.12906 | 0.41067 | 2.644 | -0.1734
390.7635 | 0.051721 | 0.002297 | 0.001058 | -0.000262 | 0.006272 | 0.038298 | -0.001536 | 0.62221 | 0.083489 | 6.3299 | 390.7635 | 6.3023 | 0.27989 | 0.12892 | -0.031925 | 0.76425 | 4.6667 | -0.18716
391.451 | 0.05605 | 0.001772 | 0.002059 | 0.001755 | 0.007141 | 0.038758 | -0.001556 | 0.6445 | 0.095135 | 6.4299 | 391.451 | 5.8779 | 0.18583 | 0.21592 | 0.18404 | 0.74886 | 4.0645 | -0.16317
392.1386 | 0.051432 | 0.00256 | 0.000775 | 0.00171 | 0.002531 | 0.02875 | 0.001008 | 0.67886 | 0.061842 | 6.5308 | 392.1386 | 8.0009 | 0.39824 | 0.12056 | 0.26601 | 0.39373 | 4.4724 | 0.15681
392.8262 | 0.050516 | 0.002571 | 0.001046 | 0.0002 | 0.004818 | 0.025811 | -0.002925 | 0.63868 | 0.063121 | 6.6326 | 392.8262 | 8.311 | 0.42299 | 0.17209 | 0.032904 | 0.79267 | 4.2465 | -0.48123
393.5138 | 0.0561 | 0.001149 | 0.0009 | 0.001714 | 0.008528 | 0.038539 | 0.001216 | 0.6578 | 0.036732 | 6.7353 | 393.5138 | 15.6378 | 0.32028 | 0.25087 | 0.47778 | 2.3772 | 10.7427 | 0.33896
394.2013 | 0.053858 | 0.001052 | 0.000535 | 0.00158 | -0.001001 | 0.023879 | -0.000835 | 0.66084 | 0.098356 | 6.8389 | 394.2013 | 5.6668 | 0.11069 | 0.056291 | 0.16624 | -0.10532 | 2.5125 | -0.087856
394.8889 | 0.05153 | 0.003662 | 0.001248 | 0.001047 | 0.007983 | 0.029445 | -0.002034 | 0.66405 | 0.098979 | 6.9434 | 394.8889 | 5.4437 | 0.38686 | 0.13184 | 0.11061 | 0.84333 | 3.1106 | -0.21487
395.5765 | 0.056034 | 0.001321 | 0.001467 | 0.002149 | 0.003332 | 0.020025 | -0.001939 | 0.72543 | 0.027391 | 7.0489 | 395.5765 | 19.8777 | 0.46862 | 0.52041 | 0.76234 | 1.182 | 7.1037 | -0.68785
396.2641 | 0.054754 | 0.002598 | 0.000364 | 0.001791 | 0.005836 | 0.017012 | -0.004624 | 0.71437 | 0.09596 | 7.1553 | 396.2641 | 5.7152 | 0.27118 | 0.037994 | 0.18694 | 0.60916 | 1.7757 | -0.48265
396.9516 | 0.054573 | 0.002744 | 0.00149 | -0.000746 | 0.00249 | 0.023042 | -0.002838 | 0.66669 | 0.099426 | 7.2626 | 396.9516 | 5.9792 | 0.30064 | 0.16325 | -0.081735 | 0.27281 | 2.5246 | -0.31094
397.6392 | 0.054507 | 0.003335 | 0.000946 | -0.00092 | 0.010893 | 0.026147 | 0.01006 | 0.66913 | 0.078106 | 7.3709 | 397.6392 | 7.6875 | 0.47036 | 0.13342 | -0.12975 | 1.5363 | 3.6877 | 1.4188
398.3268 | 0.053318 | 0.001418 | 0.000384 | 0.000343 | 0.012298 | 0.020605 | -0.003033 | 0.7074 | 0.088008 | 7.4802 | 398.3268 | 6.4062 | 0.17037 | 0.046138 | 0.041212 | 1.4776 | 2.4757 | -0.36442
399.0144 | 0.056594 | 0.001132 | 0.0006 | 0.003049 | 0.007729 | 0.033641 | 0.004429 | 0.70612 | 0.1177 | 7.5904 | 399.0144 | 5.1686 | 0.10338 | 0.054797 | 0.27846 | 0.70587 | 3.0724 | 0.40449
399.702 | 0.055 | 0.002783 | 0.001391 | 0.002314 | 0.004187 | 0.042614 | 0.000467 | 0.70265 | 0.099513 | 7.7015 | 399.702 | 6.0579 | 0.30653 | 0.15321 | 0.25487 | 0.46117 | 4.6936 | 0.051437
400.3895 | 0.053256 | 0.0029 | 0.001919 | 0.002022 | 0.003327 | 0.014651 | -0.004536 | 0.74484 | 0.087399 | 7.8136 | 400.3895 | 6.3922 | 0.34808 | 0.23033 | 0.2427 | 0.39934 | 1.7585 | -0.54445
401.0771 | 0.050764 | 0.002201 | 0.000931 | 0.00099 | 0.004511 | 0.038143 | 0.003627 | 0.7023 | 0.060348 | 7.9267 | 401.0771 | 9.4944 | 0.41165 | 0.17413 | 0.18516 | 0.8437 | 7.1339 | 0.67836
401.7647 | 0.056183 | 0.003696 | 0.002105 | 0.000603 | 0.006478 | 0.025206 | 0.001527 | 0.70874 | 0.046895 | 8.0408 | 401.7647 | 13.592 | 0.89415 | 0.50925 | 0.14588 | 1.5672 | 6.0979 | 0.36942
402.4523 | 0.05562 | 0.003456 | 0.000764 | 0.001196 | 0.00628 | 0.02067 | 0.003459 | 0.72034 | 0.057951 | 8.1558 | 402.4523 | 10.8668 | 0.67522 | 0.14927 | 0.23367 | 1.227 | 4.0384 | 0.67581
403.1398 | 0.055057 | 0.002805 | 0.00104 | 0.001448 | 0.008502 | 0.021197 | 0.001949 | 0.7219 | 0.026556 | 8.2719 | 403.1398 | 23.7565 | 1.2103 | 0.44875 | 0.6248 | 3.6685 | 9.1463 | 0.84097
403.8274 | 0.05529 | 0.002047 | 0.001203 | 0.00037 | 0.008046 | 0.02908 | 0.000849 | 0.7255 | 0.094285 | 8.3889 | 403.8274 | 6.7807 | 0.25104 | 0.14753 | 0.045376 | 0.98675 | 3.5663 | 0.10412
404.515 | 0.050559 | 0.000881 | 0.001438 | 0.001189 | 0.007981 | 0.020572 | -0.007098 | 0.74029 | 0.11211 | 8.507 | 404.515 | 5.1825 | 0.090306 | 0.1474 | 0.12188 | 0.81808 | 2.1087 | -0.72757
405.2026 | 0.056204 | -0.000373 | 0.0007 | 0.001806 | 0.005004 | 0.035406 | -0.005086 | 0.75009 | 0.11851 | 8.626 | 405.2026 | 5.454 | -0.036195 | 0.067927 | 0.17525 | 0.48558 | 3.4358 | -0.49354
405.8902 | 0.052231 | 0.003743 | 0.000468 | 0.001697 | 0.003841 | 0.014135 | 0.001881 | 0.73754 | 0.11593 | 8.746 | 405.8902 | 5.3428 | 0.38288 | 0.047872 | 0.17359 | 0.3929 | 1.4459 | 0.19241
406.5777 | 0.054667 | 0.001743 | -0.000271 | 0.001516 | 0.008819 | 0.009471 | 0.004086 | 0.74133 | -0.019312 | 8.8671 | 406.5777 | -33.858 | -1.0795 | 0.16784 | -0.93893 | -5.462 | -5.8659 | -2.5307
407.2653 | 0.055753 | 0.004066 | 0.000798 | 0.000521 | 0.010487 | 0.027631 | 0.000615 | 0.71762 | 0.17276 | 8.9892 | 407.2653 | 4.0425 | 0.29481 | 0.05786 | 0.037776 | 0.76038 | 2.0034 | 0.044592
407.9529 | 0.054781 | 0.002225 | 0.000418 | 0.001529 | 0.007797 | 0.028929 | -0.001911 | 0.71527 | 0.096008 | 9.1123 | 407.9529 | 7.2691 | 0.29524 | 0.055466 | 0.20289 | 1.0346 | 3.8387 | -0.25358
408.6405 | 0.057069 | 0.001919 | 0.001252 | 0.001037 | 0.006468 | 0.022442 | -0.001222 | 0.74094 | 0.1404 | 9.2365 | 408.6405 | 5.067 | 0.17038 | 0.11116 | 0.092073 | 0.57428 | 1.9926 | -0.1085
409.328 | 0.053235 | 0.002227 | 0.00191 | 0.00042 | 0.01287 | 0.010432 | -0.006862 | 0.74869 | 0.069232 | 9.3616 | 409.328 | 9.6147 | 0.40222 | 0.34496 | 0.075856 | 2.3244 | 1.8841 | -1.2393
410.0156 | 0.053322 | 0.00291 | 0.000291 | 0.001866 | 0.005205 | 0.014466 | 0.002927 | 0.74447 | 0.12904 | 9.4879 | 410.0156 | 5.2663 | 0.2874 | 0.02874 | 0.18429 | 0.51407 | 1.4287 | 0.28908
410.7032 | 0.054721 | 0.004101 | 0.000257 | 0.00132 | 0.003406 | 0.032883 | 0.003615 | 0.7469 | 0.11659 | 9.6152 | 410.7032 | 6.0421 | 0.45282 | 0.028377 | 0.14575 | 0.37608 | 3.6308 | 0.39915
411.3908 | 0.054251 | 0.001845 | -1e-06 | 0.002625 | 0.003915 | 0.028847 | 0.004725 | 0.75721 | 0.030232 | 9.7435 | 411.3908 | 23.0908 | 0.78528 | -0.00042563 | 1.1173 | 1.6663 | 12.2781 | 2.0111
412.0784 | 0.054008 | 0.001358 | 0.000874 | 0.001781 | 0.007391 | 0.027266 | -0.005073 | 0.75116 | 0.078631 | 9.8728 | 412.0784 | 9.0276 | 0.22699 | 0.14609 | 0.2977 | 1.2354 | 4.5576 | -0.84797
412.7659 | 0.053516 | 0.005054 | 0.000745 | 0.002157 | 0.009243 | 0.016807 | 0.00394 | 0.75731 | 0.080451 | 10.0031 | 412.7659 | 8.7865 | 0.82979 | 0.12232 | 0.35414 | 1.5176 | 2.7594 | 0.64688
413.4535 | 0.057189 | 0.0038 | 0.000676 | 0.0004 | 0.01202 | 0.02688 | -0.004254 | 0.77373 | 0.05996 | 10.1344 | 413.4535 | 12.4928 | 0.8301 | 0.14767 | 0.087379 | 2.6257 | 5.8719 | -0.92927
414.1411 | 0.053732 | 0.001982 | 0.000524 | 0.000927 | 0.007048 | 0.024024 | -0.004256 | 0.76078 | 0.077465 | 10.2667 | 414.1411 | 9.3605 | 0.34528 | 0.091285 | 0.16149 | 1.2278 | 4.1852 | -0.74143
414.8287 | 0.053344 | 0.001938 | -0.000112 | 0.001219 | 0.008597 | 0.025778 | -0.012403 | 0.74506 | 0.10313 | 10.3999 | 414.8287 | 7.2199 | 0.2623 | -0.015159 | 0.16499 | 1.1636 | 3.4889 | -1.6787
415.5162 | 0.057628 | 0.003199 | 0.000304 | 0.001998 | 0.006862 | 0.038942 | -0.004437 | 0.77679 | 0.023253 | 10.534 | 415.5162 | 33.6081 | 1.8656 | 0.17729 | 1.1652 | 4.0019 | 22.7106 | -2.5876
416.2038 | 0.051206 | 0.00407 | 0.002447 | 0.00358 | 0.005912 | 0.028965 | 0.000578 | 0.73679 | 0.071766 | 10.6691 | 416.2038 | 10.3321 | 0.82122 | 0.49374 | 0.72235 | 1.1929 | 5.8444 | 0.11663
416.8914 | 0.055092 | 0.003996 | 0.000592 | 0.001476 | 0.007013 | 0.034407 | -0.002364 | 0.79747 | 0.046905 | 10.805 | 416.8914 | 15.9139 | 1.1543 | 0.17101 | 0.42636 | 2.0258 | 9.9388 | -0.68287
417.579 | 0.050942 | 0.00259 | 0.001343 | 0.001862 | 0.007099 | 0.029726 | 0.000351 | 0.7558 | 0.075057 | 10.9417 | 417.579 | 9.8257 | 0.49956 | 0.25904 | 0.35914 | 1.3693 | 5.7336 | 0.067701
418.2665 | 0.052847 | 0.004339 | 0.00134 | 0.000126 | 0.010893 | 0.025082 | -0.005107 | 0.81872 | 0.079135 | 11.0794 | 418.2665 | 9.0371 | 0.74199 | 0.22915 | 0.021547 | 1.8628 | 4.2891 | -0.87332
418.9541 | 0.052935 | 0.003471 | 0.000739 | 0.001321 | 0.006016 | 0.032409 | -0.000353 | 0.80036 | 0.032882 | 11.2178 | 418.9541 | 22.5636 | 1.4795 | 0.315 | 0.56308 | 2.5643 | 13.8144 | -0.15047
419.6417 | 0.054765 | 0.001157 | 0.00094 | 0.001037 | 0.011351 | 0.019962 | -0.003996 | 0.85678 | 0.081916 | 11.3571 | 419.6417 | 8.862 | 0.18722 | 0.15211 | 0.16781 | 1.8368 | 3.2302 | -0.64663
420.3293 | 0.052153 | 0.002646 | 0.000861 | 0.000584 | 0.010707 | 0.033162 | -0.014892 | 0.76222 | 0.043704 | 11.4972 | 420.3293 | 17.9999 | 0.91323 | 0.29716 | 0.20156 | 3.6954 | 11.4454 | -5.1398
421.0169 | 0.05172 | 0.003405 | 0.000478 | 0.00102 | 0.013151 | 0.036901 | -0.006071 | 0.82581 | 0.020102 | 11.6381 | 421.0169 | 36.2599 | 2.3872 | 0.33512 | 0.7151 | 9.2199 | 25.8706 | -4.2563
421.7044 | 0.05364 | 0.002627 | 0.000886 | 0.002092 | 0.009764 | 0.048724 | -0.000439 | 0.85371 | 0.12157 | 11.7798 | 421.7044 | 6.0883 | 0.29817 | 0.10056 | 0.23745 | 1.1082 | 5.5303 | -0.049828
422.392 | 0.052427 | 0.002198 | 0.000381 | -0.000677 | 0.002455 | 0.039163 | -0.003214 | 0.84683 | 0.081904 | 11.9224 | 422.392 | 9.0119 | 0.37782 | 0.065492 | -0.11637 | 0.422 | 6.7319 | -0.55247
423.0796 | 0.048078 | 0.002165 | 0.000302 | 0.003584 | 0.012375 | 0.02462 | -0.002349 | 0.86412 | 0.083253 | 12.0659 | 423.0796 | 8.0636 | 0.36311 | 0.050651 | 0.60111 | 2.0755 | 4.1293 | -0.39397
423.7672 | 0.050534 | 0.002979 | 0.001174 | 0.001835 | 0.010947 | 0.028906 | -0.004889 | 0.87172 | 0.027736 | 12.2103 | 423.7672 | 25.5201 | 1.5044 | 0.59288 | 0.92669 | 5.5283 | 14.5978 | -2.469
424.4547 | 0.051676 | 0.004893 | 0.000715 | 0.001207 | 0.00724 | 0.048618 | -0.007188 | 0.85846 | 0.10295 | 12.3558 | 424.4547 | 7.2249 | 0.6841 | 0.099966 | 0.16875 | 1.0122 | 6.7974 | -1.005
425.1423 | 0.051545 | 0.001845 | 0.001513 | -0.001447 | 0.006232 | 0.055953 | -0.006305 | 0.86631 | 0.15165 | 12.5023 | 425.1423 | 4.9053 | 0.17558 | 0.14398 | -0.1377 | 0.59306 | 5.3247 | -0.60001
425.8299 | 0.051793 | 0.002317 | 0.000636 | 0.001054 | 0.007548 | 0.032239 | -0.009744 | 0.84643 | 0.04319 | 12.6499 | 425.8299 | 17.9219 | 0.80175 | 0.22008 | 0.36472 | 2.6118 | 11.1557 | -3.3717
426.5175 | 0.050731 | 0.002342 | -0.000369 | -0.000441 | 0.008767 | 0.039254 | 0.000872 | 0.87045 | 0.089291 | 12.7985 | 426.5175 | 8.3538 | 0.38565 | -0.060763 | -0.072619 | 1.4437 | 6.4639 | 0.14359
427.2051 | 0.049962 | 0.002118 | 0.000723 | 0.002027 | 0.006819 | 0.016319 | -0.003947 | 0.89748 | 0.085463 | 12.9484 | 427.2051 | 8.4343 | 0.35755 | 0.12205 | 0.34219 | 1.1512 | 2.7549 | -0.66631
427.8926 | 0.053314 | 0.001129 | 0.000769 | 0.002468 | 0.010125 | 0.025203 | -0.000367 | 0.89487 | 0.17968 | 13.0994 | 427.8926 | 4.3435 | 0.091979 | 0.06265 | 0.20107 | 0.82488 | 2.0533 | -0.029899
428.5802 | 0.059867 | 0.002601 | 0.001219 | 0.001951 | 0.008614 | 0.025602 | 0.001179 | 0.88628 | 0.097197 | 13.2516 | 428.5802 | 9.2094 | 0.40011 | 0.18752 | 0.30012 | 1.3251 | 3.9384 | 0.18137
429.2678 | 0.053449 | 0.002337 | 0.000285 | -0.000115 | 0.008036 | 0.019785 | -0.004969 | 0.87895 | 0.033874 | 13.4051 | 429.2678 | 24.0649 | 1.0522 | 0.12832 | -0.051778 | 3.6181 | 8.908 | -2.2372
429.9554 | 0.054262 | 0.001811 | 0.000354 | -0.001149 | 0.007496 | 0.029232 | -0.008171 | 0.92304 | 0.088451 | 13.5599 | 429.9554 | 9.0122 | 0.30078 | 0.058795 | -0.19083 | 1.245 | 4.855 | -1.3571
430.6429 | 0.052829 | 0.002009 | 0.001171 | 0.003527 | 0.007504 | 0.025004 | 0.001101 | 0.93592 | 0.043245 | 13.716 | 430.6429 | 17.903 | 0.68082 | 0.39684 | 1.1952 | 2.543 | 8.4735 | 0.37311
431.3305 | 0.057769 | 0.001928 | 0.001045 | -0.002362 | 0.006107 | 0.027437 | -0.003702 | 0.92165 | 0.056261 | 13.8734 | 431.3305 | 15.4562 | 0.51584 | 0.27959 | -0.63196 | 1.6339 | 7.3408 | -0.99048
432.0181 | 0.055296 | 0.003358 | 0.001452 | 0.002179 | 0.011729 | 0.031424 | 0.006238 | 0.95357 | 0.037709 | 14.0321 | 432.0181 | 21.5783 | 1.3104 | 0.56662 | 0.85032 | 4.577 | 12.2627 | 2.4343
432.7057 | 0.056603 | 0.000414 | 0.000396 | 0.001201 | 0.008145 | 0.021747 | 0.008016 | 0.9967 | 0.064959 | 14.1918 | 432.7057 | 12.4072 | 0.090748 | 0.086802 | 0.26326 | 1.7854 | 4.7669 | 1.7571
433.3933 | 0.057257 | 0.002781 | 0.000552 | 0.002328 | 0.008104 | 0.024498 | -0.002885 | 0.94167 | 0.10918 | 14.3527 | 433.3933 | 7.9934 | 0.38824 | 0.077063 | 0.325 | 1.1314 | 3.4201 | -0.40276
434.0808 | 0.055367 | 0.003824 | 0.001149 | 0.000731 | 0.007754 | 0.025023 | 0.00208 | 0.99736 | 0.082799 | 14.5145 | 434.0808 | 9.7315 | 0.67212 | 0.20195 | 0.12848 | 1.3629 | 4.3981 | 0.36559
434.7684 | 0.052088 | 0.00355 | 0.000917 | 0.002316 | 0.005551 | 0.034198 | -0.004545 | 1.027 | 0.1172 | 14.6774 | 434.7684 | 6.352 | 0.43291 | 0.11183 | 0.28243 | 0.67693 | 4.1703 | -0.55425
435.456 | 0.055015 | 0.003522 | 0.000111 | -0.00067 | 0.006227 | 0.040748 | 0.00105 | 1.0244 | 0.14052 | 14.841 | 435.456 | 5.6717 | 0.3631 | 0.011443 | -0.069073 | 0.64196 | 4.2009 | 0.10825
436.1341 | 0.053054 | 0.003598 | 0.000865 | 0.002904 | 0.006219 | 0.033927 | -0.003334 | 1.054 | 0.067516 | 15.0033 | 436.1341 | 11.1855 | 0.75857 | 0.18237 | 0.61226 | 1.3112 | 7.1529 | -0.70291
436.8123 | 0.060254 | 0.002722 | 0.001622 | -0.001425 | 0.009498 | 0.037023 | 0.003402 | 1.059 | 0.012353 | 15.1662 | 436.8123 | 69.8606 | 3.156 | 1.8806 | -1.6522 | 11.0123 | 42.9258 | 3.9444
437.4904 | 0.059029 | 0.002928 | 0.000705 | -8.1e-05 | 0.008172 | 0.021643 | 0.000197 | 1.07 | 0.051337 | 15.3298 | 437.4904 | 16.4737 | 0.81714 | 0.19675 | -0.022605 | 2.2806 | 6.0401 | 0.054978
438.1685 | 0.060612 | 0.001701 | 0.001729 | 0.001741 | 0.011685 | 0.024634 | 0.004954 | 1.0575 | 0.052067 | 15.494 | 438.1685 | 17.0566 | 0.47867 | 0.48655 | 0.48993 | 3.2882 | 6.9322 | 1.3941
438.8467 | 0.059557 | 0.003001 | 0.001005 | 0.001874 | 0.007038 | 0.01763 | -0.004338 | 1.0939 | 0.070443 | 15.6588 | 438.8467 | 12.1025 | 0.60983 | 0.20423 | 0.38081 | 1.4302 | 3.5826 | -0.88152
439.5248 | 0.060482 | 0.001658 | 0.001938 | 0.002165 | 0.01131 | 0.015922 | -0.004105 | 1.1263 | 0.019184 | 15.824 | 439.5248 | 44.2927 | 1.2142 | 1.4193 | 1.5855 | 8.2826 | 11.6601 | -3.0062
440.2029 | 0.059629 | 0.001822 | 0.000406 | 0.000564 | 0.006017 | 0.056395 | -0.008181 | 1.1242 | 0.061672 | 15.9896 | 440.2029 | 13.7515 | 0.42018 | 0.093631 | 0.13007 | 1.3876 | 13.0057 | -1.8867
440.8811 | 0.059628 | 0.003861 | 0.001198 | 0.001477 | 0.008118 | 0.038162 | 0.005627 | 1.1315 | -0.040934 | 16.1556 | 440.8811 | -20.799 | -1.3468 | -0.41788 | -0.5152 | -2.8317 | -13.3114 | -1.9628
441.5592 | 0.059945 | 0.00388 | 0.001159 | 0.004645 | 0.008817 | 0.021251 | 0.000154 | 1.15 | 0.051991 | 16.322 | 441.5592 | 16.3643 | 1.0592 | 0.31639 | 1.268 | 2.4069 | 5.8013 | 0.04204
442.2373 | 0.06138 | 0.001098 | 0.001356 | -0.001448 | 0.007219 | 0.036873 | 0.001003 | 1.229 | 0.056106 | 16.489 | 442.2373 | 14.6776 | 0.26256 | 0.32426 | -0.34626 | 1.7263 | 8.8173 | 0.23984
442.9155 | 0.055482 | -0.000508 | 0.000873 | 0.002359 | 0.008324 | 0.030536 | -0.00495 | 1.2021 | 0.05814 | 16.6566 | 442.9155 | 13.2232 | -0.12107 | 0.20806 | 0.56223 | 1.9839 | 7.2777 | -1.1797
443.5936 | 0.061985 | 0.001683 | 0.001008 | 0.001131 | 0.009828 | 0.020867 | -0.003223 | 1.1886 | 0.0724 | 16.8248 | 443.5936 | 12.1194 | 0.32906 | 0.19709 | 0.22114 | 1.9216 | 4.08 | -0.63017
444.2718 | 0.056388 | 0.003625 | 0.000807 | 0.001895 | 0.010196 | 0.032715 | -0.004883 | 1.2413 | 0.04702 | 16.9937 | 444.2718 | 16.4177 | 1.0554 | 0.23496 | 0.55174 | 2.9686 | 9.5252 | -1.4217
444.9499 | 0.061533 | 0.002767 | 0.001591 | 0.002208 | 0.00966 | 0.032015 | 0.000961 | 1.293 | 0.06601 | 17.1635 | 444.9499 | 12.3739 | 0.55643 | 0.31994 | 0.44402 | 1.9426 | 6.438 | 0.19325
445.628 | 0.057983 | 0.003498 | 0.001051 | -0.000107 | 0.01116 | 0.049871 | 0.001377 | 1.268 | 0.059532 | 17.3341 | 445.628 | 13.3148 | 0.80326 | 0.24134 | -0.024571 | 2.5627 | 11.452 | 0.3162
446.3062 | 0.056154 | 0.000987 | 0.000474 | 0.001918 | 0.013404 | 0.040354 | 0.002957 | 1.2724 | 0.10718 | 17.5057 | 446.3062 | 7.208 | 0.12669 | 0.060844 | 0.2462 | 1.7206 | 5.1799 | 0.37957
446.9843 | 0.059742 | 0.003184 | 0.001603 | 0.001766 | 0.005938 | 0.020466 | 0.001142 | 1.3388 | 0.051355 | 17.6783 | 446.9843 | 15.3606 | 0.81866 | 0.41216 | 0.45407 | 1.5268 | 5.2621 | 0.29363
447.6624 | 0.061404 | 0.004256 | 0.001559 | 0.001742 | 0.009162 | 0.035053 | 0.00545 | 1.349 | 0.031177 | 17.852 | 447.6624 | 26.0649 | 1.8066 | 0.66177 | 0.73945 | 3.8891 | 14.8794 | 2.3134
448.3406 | 0.065086 | 0.002892 | 0.001205 | 0.001336 | 0.009316 | 0.046175 | -0.008389 | 1.4088 | 0.082639 | 18.0268 | 448.3406 | 10.0781 | 0.44781 | 0.18659 | 0.20687 | 1.4425 | 7.1499 | -1.299
449.0187 | 0.060067 | 0.005658 | 0.000362 | 0.003385 | 0.01278 | 0.023243 | -0.004348 | 1.3757 | 0.096057 | 18.2029 | 449.0187 | 8.2741 | 0.77937 | 0.049864 | 0.46627 | 1.7604 | 3.2017 | -0.59892
449.6968 | 0.062729 | 0.004173 | 0.001217 | 0.002426 | 0.008349 | 0.042205 | -0.00758 | 1.4364 | 0.090485 | 18.3803 | 449.6968 | 8.8711 | 0.59014 | 0.17211 | 0.34308 | 1.1807 | 5.9686 | -1.072
450.375 | 0.061977 | 0.002505 | 0.002255 | 0.001039 | 0.012981 | 0.041212 | 0.003695 | 1.4234 | 0.048192 | 18.559 | 450.375 | 16.7677 | 0.67772 | 0.61008 | 0.2811 | 3.512 | 11.1498 | 0.99967
451.0531 | 0.066251 | 0.002983 | 0.002044 | 0.001852 | 0.007689 | 0.043496 | -0.004319 | 1.4811 | 0.11189 | 18.7391 | 451.0531 | 7.4917 | 0.33732 | 0.23114 | 0.20943 | 0.86948 | 4.9186 | -0.4884
451.7312 | 0.066023 | 0.003246 | 0.001214 | -0.000405 | 0.009293 | 0.042202 | -0.000307 | 1.4803 | 0.090291 | 18.9205 | 451.7312 | 9.3461 | 0.4595 | 0.17185 | -0.057331 | 1.3155 | 5.974 | -0.043458
452.4094 | 0.063049 | 0.005342 | 0.001644 | 0.000325 | 0.005418 | 0.050407 | 0.00531 | 1.5711 | 0.089902 | 19.1031 | 452.4094 | 8.5272 | 0.72249 | 0.22235 | 0.043955 | 0.73277 | 6.8174 | 0.71816
453.0875 | 0.065314 | 0.002853 | 0.001394 | 0.002131 | 0.011371 | 0.033522 | -0.004248 | 1.5782 | 0.092343 | 19.2869 | 453.0875 | 8.6436 | 0.37756 | 0.18448 | 0.28201 | 1.5048 | 4.4363 | -0.56218
453.7657 | 0.060585 | 0.003286 | 0.001554 | 0.001868 | 0.010076 | 0.04089 | 0.002108 | 1.5749 | 0.053279 | 19.4718 | 453.7657 | 14.0588 | 0.76252 | 0.36061 | 0.43347 | 2.3381 | 9.4885 | 0.48916
454.4438 | 0.066341 | 0.002233 | 0.002711 | -0.00078 | 0.011371 | 0.029444 | -0.003983 | 1.6347 | 0.12764 | 19.6577 | 454.4438 | 6.25 | 0.21037 | 0.2554 | -0.073484 | 1.0713 | 2.7739 | -0.37524
455.1219 | 0.064075 | 0.002839 | 0.002159 | -0.001057 | 0.012076 | 0.032673 | 0.001277 | 1.7062 | 0.076737 | 19.8445 | 455.1219 | 9.7115 | 0.43029 | 0.32723 | -0.1602 | 1.8303 | 4.9521 | 0.19355
455.8 | 0.066916 | 0.00316 | 0.001112 | 0.001357 | 0.010404 | 0.048321 | 0.001521 | 1.6787 | 0.13348 | 20.0322 | 455.8 | 5.9823 | 0.2825 | 0.099412 | 0.12132 | 0.93011 | 4.3199 | 0.13598
456.4782 | 0.061939 | 0.005468 | 0.002531 | 0.0017 | 0.005045 | 0.032176 | -0.002581 | 1.7813 | 0.11592 | 20.2207 | 456.4782 | 6.0654 | 0.53546 | 0.24785 | 0.16647 | 0.49403 | 3.1508 | -0.25275
457.1563 | 0.066814 | 0.004429 | 0.001572 | 0.001858 | 0.010903 | 0.03805 | -0.001534 | 1.8181 | 0.086809 | 20.4099 | 457.1563 | 8.6402 | 0.57275 | 0.20329 | 0.24027 | 1.4099 | 4.9205 | -0.19837
457.8344 | 0.071811 | 0.003246 | 0.001989 | 0.002752 | 0.004638 | 0.033651 | 0.00441 | 1.8564 | 0.10782 | 20.5999 | 457.8344 | 7.3906 | 0.33407 | 0.2047 | 0.28323 | 0.47733 | 3.4633 | 0.45387
458.5126 | 0.069082 | 0.002377 | 0.002026 | 0.002324 | 0.009958 | 0.034237 | -0.002974 | 1.8906 | 0.042687 | 20.7904 | 458.5126 | 17.7966 | 0.61235 | 0.52193 | 0.5987 | 2.5653 | 8.82 | -0.76615
459.1907 | 0.067668 | 0.002201 | 0.002876 | 0.002073 | 0.010895 | 0.014875 | -0.001281 | 1.9222 | 0.03986 | 20.9814 | 459.1907 | 18.5299 | 0.60271 | 0.78755 | 0.56766 | 2.9834 | 4.0733 | -0.35078
459.8689 | 0.069248 | 0.004974 | 0.00253 | 0.003048 | 0.009871 | 0.036219 | 0.000656 | 1.9936 | 0.12665 | 21.1729 | 459.8689 | 5.8069 | 0.4171 | 0.21216 | 0.25559 | 0.82775 | 3.0372 | 0.05501
460.547 | 0.068805 | 0.003634 | 0.003157 | 0.002084 | 0.011219 | 0.043276 | -0.002479 | 2.0414 | 0.056973 | 21.3648 | 460.547 | 12.6395 | 0.66757 | 0.57994 | 0.38283 | 2.0609 | 7.9498 | -0.45539
461.2251 | 0.0732 | 0.003884 | 0.002195 | -0.000416 | 0.008953 | 0.054257 | 0.002285 | 2.1095 | 0.051834 | 21.5571 | 461.2251 | 14.4316 | 0.76574 | 0.43275 | -0.082016 | 1.7651 | 10.6969 | 0.4505
461.9033 | 0.071817 | 0.005076 | 0.002847 | -0.000485 | 0.010243 | 0.047592 | -0.001624 | 2.1907 | 0.15902 | 21.75 | 461.9033 | 4.4837 | 0.31691 | 0.17775 | -0.03028 | 0.6395 | 2.9713 | -0.10139
462.5814 | 0.07226 | 0.00286 | 0.002195 | 0.002956 | 0.007728 | 0.042629 | -0.000778 | 2.2121 | 0.06529 | 21.9433 | 462.5814 | 10.9787 | 0.43453 | 0.33349 | 0.44912 | 1.1741 | 6.4768 | -0.1182
463.2595 | 0.07274 | 0.002389 | 0.001723 | 0.001299 | 0.009318 | 0.041595 | -0.004829 | 2.2593 | 0.013727 | 22.1372 | 463.2595 | 51.9206 | 1.7052 | 1.2298 | 0.9272 | 6.651 | 29.6898 | -3.4469
463.9377 | 0.075959 | 0.006079 | 0.001794 | 0.002181 | 0.008909 | 0.042247 | -0.000324 | 2.3344 | 0.0744 | 22.3318 | 463.9377 | 9.767 | 0.78165 | 0.23068 | 0.28044 | 1.1455 | 5.4322 | -0.041661
464.6158 | 0.078653 | 0.005277 | 0.00207 | 0.00173 | 0.008683 | 0.029262 | -0.001066 | 2.4652 | 0.045028 | 22.5272 | 464.6158 | 15.9623 | 1.0709 | 0.4201 | 0.3511 | 1.7622 | 5.9386 | -0.21634
465.2939 | 0.08055 | 0.004585 | 0.003254 | 0.002654 | 0.009417 | 0.048536 | -0.001465 | 2.5331 | 0.082727 | 22.7232 | 465.2939 | 8.7344 | 0.49717 | 0.35285 | 0.28779 | 1.0211 | 5.263 | -0.15886
465.9721 | 0.082047 | 0.005267 | 0.002858 | 0.001133 | 0.013009 | 0.040577 | -0.002265 | 2.6277 | 0.077652 | 22.9201 | 465.9721 | 9.2162 | 0.59163 | 0.32103 | 0.12727 | 1.4613 | 4.5579 | -0.25442
466.6502 | 0.084043 | 0.004283 | 0.002616 | 0.001375 | 0.009751 | 0.039386 | 0.005373 | 2.7001 | 0.073287 | 23.1179 | 466.6502 | 9.8183 | 0.50036 | 0.30561 | 0.16063 | 1.1392 | 4.6013 | 0.6277
467.3283 | 0.081761 | 0.005073 | 0.002041 | 0.001184 | 0.010341 | 0.031319 | 0.009186 | 2.8137 | 0.090154 | 23.3167 | 467.3283 | 7.5153 | 0.4663 | 0.1876 | 0.10883 | 0.95052 | 2.8788 | 0.84435
468.0065 | 0.086574 | 0.004947 | 0.002467 | 0.001523 | 0.008771 | 0.019777 | 0.002158 | 2.8942 | 0.062325 | 23.5164 | 468.0065 | 11.2866 | 0.64494 | 0.32162 | 0.19855 | 1.1435 | 2.5783 | 0.28134
468.6846 | 0.086798 | 0.003412 | 0.002293 | 0.004162 | 0.011161 | 0.04215 | 0.000235 | 2.992 | 0.073854 | 23.7172 | 468.6846 | 9.3161 | 0.36621 | 0.24611 | 0.44671 | 1.1979 | 4.524 | 0.025223
469.3628 | 0.091697 | 0.007233 | 0.00283 | 0.001278 | 0.011932 | 0.03608 | -0.001914 | 3.0901 | 0.065649 | 23.9192 | 469.3628 | 10.8118 | 0.85283 | 0.33368 | 0.15069 | 1.4069 | 4.2541 | -0.22568
470.0409 | 0.094981 | 0.003882 | 0.002244 | 0.000726 | 0.016852 | 0.056571 | -0.002129 | 3.1994 | 0.073207 | 24.1223 | 470.0409 | 9.7821 | 0.39981 | 0.23111 | 0.074771 | 1.7356 | 5.8263 | -0.21927
470.719 | 0.097167 | 0.005104 | 0.002642 | 0.002579 | 0.011291 | 0.054889 | -0.006702 | 3.3539 | 0.05422 | 24.3266 | 470.719 | 12.9985 | 0.68278 | 0.35343 | 0.345 | 1.5104 | 7.3427 | -0.89656
471.3972 | 0.1026 | 0.006571 | 0.00219 | 0.001445 | 0.007383 | 0.057647 | 0.00522 | 3.4497 | 0.095873 | 24.5321 | 471.3972 | 7.6106 | 0.4874 | 0.16244 | 0.10718 | 0.54763 | 4.2759 | 0.38719
472.0753 | 0.099596 | 0.003646 | 0.001466 | 0.002781 | 0.005496 | 0.035974 | -0.001823 | 3.6126 | 0.057315 | 24.7387 | 472.0753 | 11.8995 | 0.43561 | 0.17515 | 0.33227 | 0.65665 | 4.2981 | -0.21781
472.7534 | 0.10308 | 0.004675 | 0.002862 | 0.004721 | 0.016183 | 0.028965 | 0.006306 | 3.7056 | 0.06275 | 24.9464 | 472.7534 | 11.0593 | 0.50155 | 0.30705 | 0.50649 | 1.7362 | 3.1075 | 0.67653
473.4315 | 0.10803 | 0.00456 | 0.002882 | 0.000981 | 0.012753 | 0.033231 | -0.006079 | 3.8386 | 0.064004 | 25.1552 | 473.4315 | 11.0608 | 0.46689 | 0.29508 | 0.10044 | 1.3057 | 3.4024 | -0.62241
474.1097 | 0.1099 | 0.004614 | 0.002476 | 0.003135 | 0.014317 | 0.048743 | -0.004295 | 4.0349 | 0.068872 | 25.3649 | 474.1097 | 10.0311 | 0.42115 | 0.226 | 0.28615 | 1.3068 | 4.4491 | -0.39203
474.7878 | 0.11311 | 0.005959 | 0.001912 | 0.002028 | 0.014099 | 0.040961 | 0.00082 | 4.1757 | 0.075923 | 25.5755 | 474.7878 | 9.1249 | 0.48072 | 0.15424 | 0.1636 | 1.1374 | 3.3044 | 0.066151
475.466 | 0.12174 | 0.005879 | 0.002994 | 0.003122 | 0.016539 | 0.058627 | -0.000927 | 4.337 | 0.090122 | 25.787 | 475.466 | 8.0319 | 0.38787 | 0.19753 | 0.20598 | 1.0912 | 3.868 | -0.06116
476.1441 | 0.12374 | 0.003677 | 0.002553 | 0.003199 | 0.014753 | 0.056358 | -0.001102 | 4.4608 | 0.081292 | 25.9994 | 476.1441 | 8.8717 | 0.26363 | 0.18304 | 0.22936 | 1.0577 | 4.0407 | -0.07901
476.8222 | 0.12637 | 0.004369 | 0.002346 | 0.001162 | 0.013207 | 0.047742 | 0.000302 | 4.6414 | 0.070244 | 26.2125 | 476.8222 | 10.1598 | 0.35126 | 0.18861 | 0.093423 | 1.0618 | 3.8384 | 0.02428
477.5004 | 0.12964 | 0.006691 | 0.003628 | 0.003744 | 0.011654 | 0.050899 | 0.001971 | 4.8076 | 0.08522 | 26.4263 | 477.5004 | 8.3622 | 0.43158 | 0.23401 | 0.24149 | 0.75169 | 3.283 | 0.12713
478.1785 | 0.13181 | 0.004837 | 0.002669 | 0.000957 | 0.010488 | 0.058986 | -0.006042 | 4.9271 | 0.074951 | 26.6408 | 478.1785 | 9.5091 | 0.34894 | 0.19254 | 0.069038 | 0.75661 | 4.2553 | -0.43587
478.8566 | 0.13936 | 0.005604 | 0.002759 | -8e-05 | 0.012403 | 0.041989 | -0.00671 | 5.1202 | 0.076068 | 26.856 | 478.8566 | 9.6096 | 0.38641 | 0.19024 | -0.0055163 | 0.85523 | 2.8953 | -0.46268
479.5348 | 0.14509 | 0.006804 | 0.002725 | 0.002607 | 0.012941 | 0.047503 | -0.002133 | 5.2859 | 0.070849 | 27.0717 | 479.5348 | 10.4885 | 0.49185 | 0.19698 | 0.18845 | 0.93548 | 3.4339 | -0.15419
480.2129 | 0.14664 | 0.004204 | 0.003871 | 0.00381 | 0.010198 | 0.040482 | 0.001949 | 5.4783 | 0.086828 | 27.2879 | 480.2129 | 8.4123 | 0.24117 | 0.22207 | 0.21857 | 0.58503 | 2.3223 | 0.11181
480.8911 | 0.15556 | 0.005504 | 0.002691 | 0.002731 | 0.011827 | 0.053844 | 0.00027 | 5.6329 | 0.080683 | 27.5047 | 480.8911 | 9.4144 | 0.3331 | 0.16286 | 0.16528 | 0.71576 | 3.2586 | 0.01634
481.5692 | 0.15788 | 0.005448 | 0.003668 | 0.000744 | 0.01541 | 0.03271 | -0.000119 | 5.8477 | 0.0729 | 27.7219 | 481.5692 | 10.2666 | 0.35428 | 0.23852 | 0.048381 | 1.0021 | 2.1271 | -0.0077384
482.2473 | 0.16329 | 0.004948 | 0.002932 | 0.000743 | 0.016723 | 0.041671 | 0.00695 | 6.0566 | 0.094977 | 27.9397 | 482.2473 | 7.931 | 0.24033 | 0.14241 | 0.036088 | 0.81225 | 2.024 | 0.33757
482.9254 | 0.16889 | 0.004053 | 0.003194 | 0.002599 | 0.014357 | 0.051297 | -0.003158 | 6.1979 | 0.10239 | 28.158 | 482.9254 | 7.4936 | 0.17983 | 0.14172 | 0.11532 | 0.63702 | 2.276 | -0.14012
483.6036 | 0.17588 | 0.005365 | 0.002279 | 0.001779 | 0.015183 | 0.046161 | 0.001561 | 6.4304 | 0.085506 | 28.3768 | 483.6036 | 9.077 | 0.27688 | 0.11762 | 0.091812 | 0.78358 | 2.3823 | 0.080562
484.2817 | 0.17615 | 0.005456 | 0.003091 | 0.002692 | 0.013328 | 0.051137 | -0.006133 | 6.6547 | 0.086657 | 28.5962 | 484.2817 | 8.7348 | 0.27055 | 0.15328 | 0.13349 | 0.66091 | 2.5358 | -0.30412
484.9598 | 0.18713 | 0.007731 | 0.003425 | 0.000449 | 0.014888 | 0.057031 | -0.000644 | 6.84 | 0.090015 | 28.8161 | 484.9598 | 8.7582 | 0.36183 | 0.1603 | 0.021014 | 0.69679 | 2.6692 | -0.03014
485.638 | 0.19764 | 0.004527 | 0.004191 | 0.00163 | 0.016676 | 0.055254 | -0.000599 | 7.0621 | 0.074543 | 29.0367 | 485.638 | 10.9012 | 0.2497 | 0.23117 | 0.089907 | 0.91981 | 3.0477 | -0.033039
486.3161 | 0.20266 | 0.006495 | 0.003558 | 0.002669 | 0.016393 | 0.068549 | 0.004222 | 7.3342 | 0.10055 | 29.2578 | 486.3161 | 8.0408 | 0.25769 | 0.14116 | 0.10589 | 0.6504 | 2.7197 | 0.16751
486.9943 | 0.20933 | 0.007462 | 0.002424 | 0.001523 | 0.015178 | 0.05451 | 0.001345 | 7.4955 | 0.06831 | 29.4795 | 486.9943 | 12.0525 | 0.42963 | 0.13956 | 0.087687 | 0.87388 | 3.1384 | 0.077439
487.6724 | 0.21522 | 0.007437 | 0.003571 | 0.001349 | 0.013424 | 0.058064 | -0.005856 | 7.807 | 0.086761 | 29.7019 | 487.6724 | 9.4374 | 0.32611 | 0.15659 | 0.059154 | 0.58865 | 2.5461 | -0.25679
488.3505 | 0.22379 | 0.007752 | 0.003977 | 0.00208 | 0.014199 | 0.045457 | -0.003583 | 7.9641 | 0.078891 | 29.9249 | 488.3505 | 10.6586 | 0.36922 | 0.18942 | 0.099068 | 0.67628 | 2.1651 | -0.17065
489.0287 | 0.24055 | 0.007048 | 0.002829 | -0.001809 | 0.016876 | 0.067306 | 0.001804 | 8.2387 | 0.092115 | 30.1485 | 489.0287 | 9.5563 | 0.27999 | 0.11239 | -0.071865 | 0.67042 | 2.6738 | 0.071666
489.7068 | 0.24081 | 0.00838 | 0.0041 | 0.000876 | 0.017886 | 0.04711 | -0.003741 | 8.5237 | 0.087186 | 30.3728 | 489.7068 | 9.8419 | 0.34249 | 0.16757 | 0.035803 | 0.73101 | 1.9254 | -0.1529
490.3849 | 0.25192 | 0.006648 | 0.003672 | -6.6e-05 | 0.017611 | 0.054506 | -0.003622 | 8.8172 | 0.090017 | 30.5978 | 490.3849 | 9.7115 | 0.25628 | 0.14156 | -0.0025443 | 0.67891 | 2.1012 | -0.13963
491.0631 | 0.25981 | 0.007968 | 0.004001 | 0.002746 | 0.016433 | 0.04676 | -0.00232 | 9.0944 | 0.094538 | 30.8234 | 491.0631 | 9.3144 | 0.28566 | 0.14344 | 0.098446 | 0.58914 | 1.6764 | -0.083174
491.7412 | 0.26725 | 0.006417 | 0.004082 | 0.004582 | 0.017685 | 0.071284 | -0.004359 | 9.3481 | 0.084496 | 31.0497 | 491.7412 | 10.5053 | 0.25225 | 0.16046 | 0.18012 | 0.69519 | 2.8021 | -0.17135
492.4193 | 0.28057 | 0.008323 | 0.00391 | 0.000904 | 0.012042 | 0.078047 | 0.002504 | 9.6685 | 0.096824 | 31.2767 | 492.4193 | 9.3738 | 0.27807 | 0.13063 | 0.030203 | 0.40232 | 2.6075 | 0.083659
493.0975 | 0.28586 | 0.00766 | 0.003623 | 0.00218 | 0.012774 | 0.051751 | -0.003564 | 9.9073 | 0.093422 | 31.5044 | 493.0975 | 9.7301 | 0.26073 | 0.12332 | 0.074203 | 0.4348 | 1.7615 | -0.12131
493.7756 | 0.29959 | 0.006984 | 0.003624 | 0.000842 | 0.017967 | 0.066736 | 0.000751 | 10.2251 | 0.098529 | 31.7327 | 493.7756 | 9.4364 | 0.21998 | 0.11415 | 0.026521 | 0.56592 | 2.102 | 0.023655
494.4537 | 0.31191 | 0.007061 | 0.004149 | 0.001021 | 0.016902 | 0.072688 | -0.00109 | 10.531 | 0.09075 | 31.9617 | 494.4537 | 10.4316 | 0.23615 | 0.13876 | 0.034146 | 0.56527 | 2.431 | -0.036454
495.1319 | 0.32137 | 0.007036 | 0.003792 | 0.000755 | 0.018764 | 0.059659 | -0.003188 | 10.7701 | 0.097198 | 32.1914 | 495.1319 | 9.8827 | 0.21637 | 0.11661 | 0.023217 | 0.57702 | 1.8346 | -0.098036
495.81 | 0.32792 | 0.007987 | 0.004444 | 0.002278 | 0.017566 | 0.056187 | 0.000837 | 11.1859 | 0.090731 | 32.4218 | 495.81 | 10.4754 | 0.25515 | 0.14197 | 0.072772 | 0.56115 | 1.7949 | 0.026738
496.4882 | 0.34196 | 0.007567 | 0.003093 | 0.001171 | 0.017893 | 0.062798 | 0.001938 | 11.4756 | 0.092601 | 32.6528 | 496.4882 | 10.5076 | 0.23252 | 0.095041 | 0.035982 | 0.54981 | 1.9296 | 0.05955
497.1663 | 0.35161 | 0.009036 | 0.003936 | 0.002635 | 0.017694 | 0.051049 | -0.001689 | 11.7933 | 0.094859 | 32.8845 | 497.1663 | 10.3357 | 0.26562 | 0.1157 | 0.077457 | 0.52012 | 1.5006 | -0.049649
497.8444 | 0.3598 | 0.008912 | 0.003992 | 0.002978 | 0.017616 | 0.045093 | -0.003248 | 12.1386 | 0.086773 | 33.1169 | 497.8444 | 11.3126 | 0.2802 | 0.12551 | 0.093631 | 0.55386 | 1.4178 | -0.10212
498.5226 | 0.38085 | 0.010652 | 0.004085 | 0.004041 | 0.013055 | 0.050924 | -0.004345 | 12.392 | 0.093707 | 33.3499 | 498.5226 | 10.938 | 0.30592 | 0.11732 | 0.11606 | 0.37494 | 1.4625 | -0.12479
499.2007 | 0.38524 | 0.009554 | 0.003892 | 0.002991 | 0.019332 | 0.061064 | -0.003422 | 12.8444 | 0.094899 | 33.5836 | 499.2007 | 10.6141 | 0.26323 | 0.10723 | 0.082408 | 0.53263 | 1.6824 | -0.094283
499.8788 | 0.39272 | 0.011583 | 0.003535 | 0.002269 | 0.015895 | 0.070316 | -0.001402 | 13.2135 | 0.094798 | 33.818 | 499.8788 | 10.6027 | 0.31272 | 0.095438 | 0.061258 | 0.42913 | 1.8984 | -0.037851
500.5569 | 0.41486 | 0.010827 | 0.004409 | 0.002137 | 0.016417 | 0.080432 | -0.003741 | 13.5355 | 0.094913 | 34.0531 | 500.5569 | 10.9967 | 0.28699 | 0.11687 | 0.056645 | 0.43516 | 2.132 | -0.099162
501.2351 | 0.42584 | 0.011121 | 0.003926 | 0.004578 | 0.019239 | 0.077313 | -0.004297 | 13.8908 | 0.098992 | 34.2889 | 501.2351 | 10.6187 | 0.27731 | 0.097898 | 0.11416 | 0.47974 | 1.9279 | -0.10715
501.9132 | 0.43107 | 0.012882 | 0.004952 | 0.001887 | 0.020141 | 0.071838 | -0.000321 | 14.3112 | 0.10327 | 34.5253 | 501.9132 | 10.0697 | 0.30092 | 0.11568 | 0.04408 | 0.47049 | 1.6781 | -0.0074984
502.5914 | 0.45464 | 0.014482 | 0.004417 | 0.002817 | 0.014908 | 0.066955 | -0.001994 | 14.684 | 0.090972 | 34.7624 | 502.5914 | 11.8311 | 0.37687 | 0.11494 | 0.073307 | 0.38795 | 1.7424 | -0.05189
503.2695 | 0.46835 | 0.014189 | 0.004633 | 0.003085 | 0.01845 | 0.084867 | -0.002008 | 15.1492 | 0.10091 | 35.0002 | 503.2695 | 10.7228 | 0.32486 | 0.10607 | 0.070631 | 0.42241 | 1.943 | -0.045973
503.9476 | 0.47562 | 0.014406 | 0.004313 | 0.001429 | 0.019669 | 0.062634 | -0.008515 | 15.5446 | 0.09986 | 35.2388 | 503.9476 | 10.7971 | 0.32703 | 0.09791 | 0.03244 | 0.44651 | 1.4219 | -0.1933
504.6258 | 0.49615 | 0.019206 | 0.004635 | 0.004365 | 0.018814 | 0.064821 | -0.006803 | 15.9874 | 0.099625 | 35.478 | 504.6258 | 11.0517 | 0.42781 | 0.10324 | 0.09723 | 0.41908 | 1.4439 | -0.15154
505.3039 | 0.50772 | 0.01996 | 0.004693 | 0.002808 | 0.022129 | 0.069872 | -0.006209 | 16.4295 | 0.098152 | 35.718 | 505.3039 | 11.2457 | 0.4421 | 0.10395 | 0.062195 | 0.49014 | 1.5476 | -0.13753
505.982 | 0.51833 | 0.023364 | 0.004405 | 0.003282 | 0.01713 | 0.065243 | -0.004542 | 16.857 | 0.099527 | 35.9586 | 505.982 | 11.1094 | 0.50076 | 0.094413 | 0.070343 | 0.36715 | 1.3984 | -0.097349
506.6602 | 0.532 | 0.027157 | 0.00566 | 0.003486 | 0.023348 | 0.084328 | -0.00884 | 17.2737 | 0.09298 | 36.2001 | 506.6602 | 11.9908 | 0.61209 | 0.12757 | 0.078571 | 0.52624 | 1.9007 | -0.19924
507.3383 | 0.54796 | 0.031201 | 0.006314 | 0.007043 | 0.017302 | 0.081165 | -0.007089 | 17.7336 | 0.10096 | 36.4422 | 507.3383 | 11.1538 | 0.6351 | 0.12852 | 0.14336 | 0.35218 | 1.6521 | -0.1443
508.0164 | 0.56353 | 0.034045 | 0.006254 | 0.003476 | 0.019535 | 0.079519 | -0.005765 | 18.245 | 0.10053 | 36.6851 | 508.0164 | 11.2708 | 0.68091 | 0.12508 | 0.069521 | 0.39071 | 1.5904 | -0.1153
508.6946 | 0.58035 | 0.039628 | 0.005824 | 0.005292 | 0.022769 | 0.069819 | -0.005061 | 18.6488 | 0.098496 | 36.9288 | 508.6946 | 11.6678 | 0.79671 | 0.11709 | 0.10639 | 0.45776 | 1.4037 | -0.10175
509.3727 | 0.59797 | 0.045497 | 0.006799 | 0.005909 | 0.025718 | 0.066195 | -0.005879 | 19.1565 | 0.097089 | 37.1732 | 509.3727 | 11.9515 | 0.90934 | 0.13589 | 0.1181 | 0.51402 | 1.323 | -0.1175
510.0508 | 0.60842 | 0.051773 | 0.007331 | 0.00626 | 0.02227 | 0.070431 | -0.008476 | 19.6017 | 0.10042 | 37.4184 | 510.0508 | 11.5654 | 0.98414 | 0.13935 | 0.119 | 0.42333 | 1.3388 | -0.16112
510.729 | 0.63056 | 0.058554 | 0.00796 | 0.010747 | 0.022494 | 0.089318 | 0.002339 | 20.1161 | 0.093531 | 37.6644 | 510.729 | 12.6229 | 1.1722 | 0.15935 | 0.21514 | 0.4503 | 1.788 | 0.046824
511.4071 | 0.64366 | 0.06656 | 0.00921 | 0.01062 | 0.023333 | 0.084927 | 0.002539 | 20.4996 | 0.097202 | 37.9111 | 511.4071 | 12.2462 | 1.2664 | 0.17523 | 0.20206 | 0.44393 | 1.6158 | 0.048307
512.0853 | 0.66067 | 0.073169 | 0.009096 | 0.012342 | 0.026199 | 0.069748 | -0.003339 | 21.0403 | 0.097793 | 38.1585 | 512.0853 | 12.2522 | 1.3569 | 0.16869 | 0.22888 | 0.48587 | 1.2935 | -0.061922
512.7634 | 0.67593 | 0.083126 | 0.011979 | 0.014938 | 0.022943 | 0.088046 | 0.005277 | 21.4639 | 0.10123 | 38.4066 | 512.7634 | 11.9478 | 1.4693 | 0.21174 | 0.26404 | 0.40554 | 1.5563 | 0.093276
513.4415 | 0.69343 | 0.093946 | 0.012236 | 0.014138 | 0.019829 | 0.084841 | -0.004985 | 22.0634 | 0.099393 | 38.6554 | 513.4415 | 12.2232 | 1.656 | 0.21569 | 0.24921 | 0.34953 | 1.4955 | -0.087871
514.1196 | 0.6974 | 0.10453 | 0.013132 | 0.012711 | 0.020684 | 0.083947 | 0.001411 | 22.5846 | 0.10043 | 38.9048 | 514.1196 | 11.9623 | 1.7929 | 0.22525 | 0.21803 | 0.35479 | 1.4399 | 0.024202
514.7978 | 0.65638 | 0.11684 | 0.015209 | 0.016378 | 0.020398 | 0.084168 | -0.009729 | 23.0888 | 0.098261 | 39.1548 | 514.7978 | 11.328 | 2.0165 | 0.26248 | 0.28266 | 0.35204 | 1.4526 | -0.16791
515.476 | 0.7272 | 0.13031 | 0.017308 | 0.017091 | 0.030187 | 0.071865 | 4e-06 | 23.6433 | 0.096108 | 39.4054 | 515.476 | 12.6108 | 2.2598 | 0.30015 | 0.29638 | 0.52349 | 1.2462 | 6.9366e-05
516.1541 | 0.73889 | 0.1443 | 0.018739 | 0.016946 | 0.032557 | 0.096358 | -0.003135 | 24.2619 | 0.098307 | 39.6566 | 516.1541 | 12.2853 | 2.3993 | 0.31157 | 0.28175 | 0.54131 | 1.6021 | -0.052124
516.8322 | 0.76204 | 0.1617 | 0.022145 | 0.020072 | 0.028134 | 0.081665 | 0.001067 | 24.8291 | 0.10129 | 39.9082 | 516.8322 | 12.0918 | 2.5658 | 0.35139 | 0.3185 | 0.44642 | 1.2958 | 0.016931
517.5103 | 0.81516 | 0.17542 | 0.024472 | 0.022117 | 0.027753 | 0.098021 | -0.002181 | 25.4307 | 0.098987 | 40.1604 | 517.5103 | 13.0049 | 2.7986 | 0.39042 | 0.35285 | 0.44276 | 1.5638 | -0.034795
518.1885 | 0.84122 | 0.19252 | 0.027834 | 0.021482 | 0.023734 | 0.082499 | -0.002732 | 26.0408 | 0.10034 | 40.413 | 518.1885 | 13.0101 | 2.9775 | 0.43048 | 0.33224 | 0.36707 | 1.2759 | -0.042253
518.8666 | 0.86899 | 0.20952 | 0.030701 | 0.022771 | 0.024869 | 0.087273 | -0.004273 | 26.6403 | 0.098841 | 40.6661 | 518.8666 | 13.4206 | 3.2358 | 0.47414 | 0.35167 | 0.38407 | 1.3478 | -0.065991
519.5447 | 0.89114 | 0.22546 | 0.033956 | 0.023418 | 0.023261 | 0.082915 | -0.002756 | 27.2184 | 0.097129 | 40.9196 | 519.5447 | 13.7932 | 3.4897 | 0.52558 | 0.36247 | 0.36004 | 1.2834 | -0.042658
520.2229 | 0.91251 | 0.24152 | 0.038214 | 0.026341 | 0.027978 | 0.090419 | -0.013627 | 27.8931 | 0.099568 | 41.1735 | 520.2229 | 13.5282 | 3.5806 | 0.56653 | 0.39051 | 0.41478 | 1.3405 | -0.20202
520.901 | 0.93827 | 0.26372 | 0.041772 | 0.027225 | 0.031657 | 0.090603 | 9.4e-05 | 28.473 | 0.098067 | 41.4277 | 520.901 | 13.9208 | 3.9128 | 0.61976 | 0.40393 | 0.46969 | 1.3443 | 0.0013947
521.5792 | 0.96157 | 0.28431 | 0.047363 | 0.028188 | 0.028703 | 0.096468 | -0.006039 | 29.0846 | 0.099065 | 41.6823 | 521.5792 | 13.9107 | 4.113 | 0.68518 | 0.40779 | 0.41524 | 1.3956 | -0.087364
522.2573 | 0.99279 | 0.30417 | 0.052613 | 0.028967 | 0.027788 | 0.085361 | -0.005938 | 29.7393 | 0.10085 | 41.9371 | 522.2573 | 13.8815 | 4.2529 | 0.73565 | 0.40503 | 0.38854 | 1.1935 | -0.083027
522.9354 | 1.0197 | 0.32634 | 0.05708 | 0.02935 | 0.028134 | 0.075383 | 0.000737 | 30.5298 | 0.10148 | 42.1923 | 522.9354 | 13.8875 | 4.4445 | 0.77738 | 0.39972 | 0.38316 | 1.0267 | 0.010037
523.6135 | 1.0451 | 0.34963 | 0.063929 | 0.031712 | 0.028844 | 0.092602 | -0.006771 | 31.108 | 0.099754 | 42.4477 | 523.6135 | 14.2957 | 4.7826 | 0.87448 | 0.43379 | 0.39456 | 1.2667 | -0.09262
524.2917 | 1.0823 | 0.37524 | 0.069842 | 0.030935 | 0.030205 | 0.10866 | 0.005756 | 31.7271 | 0.096701 | 42.7033 | 524.2917 | 15.0649 | 5.2229 | 0.97211 | 0.43058 | 0.42042 | 1.5123 | 0.080116
524.9698 | 1.1089 | 0.39724 | 0.076485 | 0.033269 | 0.03579 | 0.097485 | 0.000383 | 32.5067 | 0.102 | 42.9591 | 524.9698 | 14.3685 | 5.147 | 0.99101 | 0.43106 | 0.46373 | 1.2631 | 0.0049625
525.6479 | 1.1453 | 0.42653 | 0.084791 | 0.03338 | 0.031941 | 0.095315 | 0.001766 | 33.2179 | 0.1004 | 43.2151 | 525.6479 | 14.8409 | 5.5268 | 1.0987 | 0.43252 | 0.41388 | 1.2351 | 0.022883
526.3261 | 1.1762 | 0.45075 | 0.091506 | 0.034925 | 0.028575 | 0.094268 | -0.00099 | 33.9717 | 0.09699 | 43.4712 | 526.3261 | 15.518 | 5.9469 | 1.2073 | 0.46078 | 0.377 | 1.2437 | -0.013062
527.0042 | 1.2068 | 0.47543 | 0.099788 | 0.033407 | 0.034468 | 0.1068 | 0.002067 | 34.6153 | 0.099409 | 43.7274 | 527.0042 | 15.3354 | 6.0416 | 1.2681 | 0.42452 | 0.438 | 1.3571 | 0.026266
527.6824 | 1.2503 | 0.50546 | 0.10893 | 0.034413 | 0.031115 | 0.10004 | -0.004899 | 35.4039 | 0.10003 | 43.9837 | 527.6824 | 15.5288 | 6.2779 | 1.353 | 0.42742 | 0.38646 | 1.2425 | -0.060847
528.3605 | 1.2801 | 0.5352 | 0.11926 | 0.036505 | 0.031666 | 0.12557 | -0.002174 | 36.1568 | 0.097242 | 44.2401 | 528.3605 | 16.107 | 6.7343 | 1.5006 | 0.45933 | 0.39844 | 1.58 | -0.027355
529.0386 | 1.3202 | 0.5625 | 0.13003 | 0.035549 | 0.03498 | 0.10689 | -0.005753 | 36.9118 | 0.098282 | 44.4965 | 529.0386 | 16.1925 | 6.8993 | 1.5949 | 0.43603 | 0.42905 | 1.3111 | -0.070564
529.7167 | 1.3473 | 0.59259 | 0.1406 | 0.033802 | 0.033941 | 0.11146 | -0.011699 | 37.7485 | 0.097562 | 44.7529 | 529.7167 | 16.3726 | 7.201 | 1.7085 | 0.41076 | 0.41244 | 1.3544 | -0.14216
530.3949 | 1.3876 | 0.62459 | 0.15182 | 0.032335 | 0.034056 | 0.10244 | 0.007494 | 38.4842 | 0.10153 | 45.0093 | 530.3949 | 15.9843 | 7.1949 | 1.7488 | 0.37248 | 0.3923 | 1.18 | 0.086326
531.0731 | 1.4255 | 0.65704 | 0.16414 | 0.0338 | 0.029354 | 0.097963 | 0.001765 | 39.3136 | 0.099406 | 45.2657 | 531.0731 | 16.5111 | 7.6104 | 1.9012 | 0.3915 | 0.34 | 1.1347 | 0.020444
531.7512 | 1.4708 | 0.68983 | 0.17662 | 0.036493 | 0.036544 | 0.098581 | -0.005689 | 40.0632 | 0.099534 | 45.5222 | 531.7512 | 16.7904 | 7.8749 | 2.0163 | 0.41659 | 0.41718 | 1.1254 | -0.064944
532.4293 | 1.5062 | 0.72313 | 0.19097 | 0.032505 | 0.036371 | 0.11335 | -0.003268 | 40.9948 | 0.097572 | 45.7788 | 532.4293 | 17.238 | 8.2761 | 2.1856 | 0.37202 | 0.41626 | 1.2973 | -0.037402
533.1074 | 1.5549 | 0.75706 | 0.2053 | 0.033402 | 0.038487 | 0.098043 | -0.001407 | 41.6943 | 0.097567 | 46.0357 | 533.1074 | 17.5958 | 8.5674 | 2.3233 | 0.378 | 0.43554 | 1.1095 | -0.015922
533.7856 | 1.5954 | 0.79678 | 0.21974 | 0.032519 | 0.037644 | 0.12012 | -0.002222 | 42.5153 | 0.098676 | 46.293 | 533.7856 | 17.6047 | 8.7922 | 2.4247 | 0.35884 | 0.41539 | 1.3255 | -0.024519
534.4637 | 1.6397 | 0.83039 | 0.23528 | 0.032135 | 0.040458 | 0.11382 | 0.001468 | 43.4262 | 0.099362 | 46.5507 | 534.4637 | 17.6892 | 8.9585 | 2.5383 | 0.34668 | 0.43647 | 1.2279 | 0.015837
535.1418 | 1.6814 | 0.86858 | 0.2514 | 0.030007 | 0.038997 | 0.12055 | 0.002811 | 44.2622 | 0.099819 | 46.8088 | 535.1418 | 17.8139 | 9.2022 | 2.6635 | 0.31791 | 0.41316 | 1.2771 | 0.029781
535.8199 | 1.7352 | 0.90595 | 0.26792 | 0.029895 | 0.041697 | 0.11252 | -0.004989 | 45.248 | 0.10059 | 47.0675 | 535.8199 | 17.9442 | 9.3685 | 2.7706 | 0.30915 | 0.43119 | 1.1635 | -0.051592
536.4981 | 1.7812 | 0.94643 | 0.28716 | 0.029562 | 0.039887 | 0.10959 | 9.7e-05 | 46.0614 | 0.10069 | 47.327 | 536.4981 | 18.1769 | 9.6581 | 2.9304 | 0.30167 | 0.40704 | 1.1183 | 0.00098986
537.1763 | 1.8333 | 0.98423 | 0.30524 | 0.028546 | 0.039422 | 0.10216 | -4.5e-05 | 47.1027 | 0.099977 | 47.5871 | 537.1763 | 18.5253 | 9.9458 | 3.0845 | 0.28846 | 0.39837 | 1.0323 | -0.00045473
537.8544 | 1.8839 | 1.0248 | 0.32497 | 0.027438 | 0.041451 | 0.11398 | -0.000407 | 48.062 | 0.098879 | 47.848 | 537.8544 | 18.968 | 10.3178 | 3.2719 | 0.27626 | 0.41734 | 1.1475 | -0.0040978
538.5325 | 1.9397 | 1.063 | 0.3452 | 0.027296 | 0.043041 | 0.11729 | -0.007126 | 49.0454 | 0.10023 | 48.1099 | 538.5325 | 18.9837 | 10.4036 | 3.3785 | 0.26715 | 0.42124 | 1.1479 | -0.069742
539.2106 | 1.9872 | 1.1082 | 0.36505 | 0.027149 | 0.038605 | 0.11602 | -0.005482 | 49.9621 | 0.10004 | 48.3727 | 539.2106 | 19.2316 | 10.7245 | 3.5329 | 0.26274 | 0.37361 | 1.1229 | -0.053054
539.8888 | 2.0386 | 1.1513 | 0.38753 | 0.025042 | 0.045379 | 0.11701 | -0.007032 | 51.0404 | 0.10051 | 48.6366 | 539.8888 | 19.3268 | 10.9145 | 3.6739 | 0.23741 | 0.43021 | 1.1093 | -0.066665
540.5669 | 2.0836 | 1.1876 | 0.40936 | 0.024944 | 0.041328 | 0.12138 | 0.004757 | 52.0514 | 0.099844 | 48.9016 | 540.5669 | 19.6058 | 11.175 | 3.852 | 0.23471 | 0.38888 | 1.1422 | 0.044761
541.2451 | 2.1468 | 1.2324 | 0.43111 | 0.025742 | 0.043163 | 0.13142 | 0.008158 | 53.1223 | 0.099286 | 49.1677 | 541.2451 | 20.0127 | 11.4886 | 4.0189 | 0.23997 | 0.40237 | 1.2251 | 0.07605
541.9232 | 2.1916 | 1.2759 | 0.45524 | 0.022952 | 0.049728 | 0.12555 | 0.002515 | 54.279 | 0.10169 | 49.4348 | 541.9232 | 19.6282 | 11.427 | 4.0771 | 0.20556 | 0.44536 | 1.1244 | 0.022524
542.6013 | 2.2421 | 1.3239 | 0.47999 | 0.020506 | 0.043784 | 0.13438 | -0.001725 | 55.4413 | 0.10086 | 49.7028 | 542.6013 | 19.9297 | 11.7676 | 4.2666 | 0.18227 | 0.38919 | 1.1945 | -0.015333
543.2795 | 2.3095 | 1.3716 | 0.50521 | 0.022007 | 0.0503 | 0.1143 | -0.004114 | 56.6246 | 0.10372 | 49.9716 | 543.2795 | 19.6517 | 11.6712 | 4.2988 | 0.18726 | 0.428 | 0.97253 | -0.035006
543.9576 | 2.3646 | 1.4146 | 0.53077 | 0.020333 | 0.050773 | 0.11912 | -0.002623 | 57.7274 | 0.10357 | 50.2411 | 543.9576 | 19.8707 | 11.8877 | 4.4603 | 0.17087 | 0.42667 | 1.001 | -0.022042
544.6357 | 2.4306 | 1.4592 | 0.55947 | 0.017855 | 0.049669 | 0.14212 | -0.005891 | 59.046 | 0.10437 | 50.5113 | 544.6357 | 19.9217 | 11.9598 | 4.5855 | 0.14634 | 0.4071 | 1.1648 | -0.048284
545.3138 | 2.4796 | 1.5129 | 0.58713 | 0.018463 | 0.051836 | 0.12998 | -0.002834 | 60.322 | 0.10386 | 50.782 | 545.3138 | 20.099 | 12.2632 | 4.7591 | 0.14965 | 0.42016 | 1.0535 | -0.022971
545.992 | 2.5509 | 1.5606 | 0.61592 | 0.016358 | 0.049993 | 0.15125 | 0.004032 | 61.5598 | 0.10367 | 51.0531 | 545.992 | 20.4051 | 12.4838 | 4.9269 | 0.13085 | 0.39991 | 1.2099 | 0.032253
546.6612 | 2.6105 | 1.614 | 0.64642 | 0.016806 | 0.053964 | 0.13395 | -0.002736 | 62.9402 | 0.10509 | 51.321 | 546.6612 | 20.2555 | 12.5235 | 5.0157 | 0.1304 | 0.41871 | 1.0393 | -0.021229
547.3304 | 2.6817 | 1.667 | 0.67935 | 0.017192 | 0.053821 | 0.14382 | 0.002708 | 64.2197 | 0.10223 | 51.5892 | 547.3304 | 21.0736 | 13.0994 | 5.3385 | 0.1351 | 0.42294 | 1.1302 | 0.02128
547.9995 | 2.7534 | 1.7258 | 0.71185 | 0.015168 | 0.054558 | 0.14626 | -0.007727 | 65.7331 | 0.10201 | 51.8576 | 547.9995 | 21.2948 | 13.347 | 5.5055 | 0.11731 | 0.42195 | 1.1312 | -0.059761
548.6687 | 2.8357 | 1.7826 | 0.74877 | 0.015883 | 0.054783 | 0.1481 | 0.001567 | 67.1669 | 0.10287 | 52.126 | 548.6687 | 21.3936 | 13.4485 | 5.6491 | 0.11983 | 0.41331 | 1.1174 | 0.011822
549.3379 | 2.9081 | 1.8458 | 0.7854 | 0.016558 | 0.056532 | 0.13516 | 0.001345 | 68.8191 | 0.10115 | 52.3945 | 549.3379 | 21.8897 | 13.8935 | 5.9118 | 0.12463 | 0.42552 | 1.0174 | 0.010124
550.0071 | 2.9949 | 1.9136 | 0.8242 | 0.013231 | 0.059006 | 0.12954 | 0.004774 | 70.4009 | 0.10136 | 52.6628 | 550.0071 | 22.1034 | 14.1232 | 6.0829 | 0.097649 | 0.43548 | 0.95603 | 0.035234
550.6763 | 3.0861 | 1.9816 | 0.86841 | 0.014765 | 0.061168 | 0.14823 | 0.00256 | 72.1589 | 0.10097 | 52.9311 | 550.6763 | 22.421 | 14.3968 | 6.3092 | 0.10727 | 0.4444 | 1.0769 | 0.018599
551.3454 | 3.1749 | 2.0592 | 0.91183 | 0.013894 | 0.066421 | 0.14532 | -0.003062 | 74.1935 | 0.10098 | 53.1991 | 551.3454 | 22.5439 | 14.6214 | 6.4746 | 0.098657 | 0.47164 | 1.0318 | -0.021742
552.0146 | 3.2861 | 2.1406 | 0.95751 | 0.01225 | 0.065704 | 0.14856 | 0.00015 | 76.086 | 0.10189 | 53.4671 | 552.0146 | 22.6644 | 14.7637 | 6.6039 | 0.084488 | 0.45316 | 1.0246 | 0.0010345
552.6838 | 3.3947 | 2.2232 | 1.0093 | 0.013852 | 0.070878 | 0.14606 | 0.000412 | 78.3881 | 0.10058 | 53.7351 | 552.6838 | 23.1353 | 15.1514 | 6.8787 | 0.094404 | 0.48305 | 0.9954 | 0.0028079
553.353 | 3.5035 | 2.314 | 1.0594 | 0.0141 | 0.070602 | 0.15654 | -0.007946 | 80.535 | 0.10061 | 54.003 | 553.353 | 23.351 | 15.4233 | 7.0613 | 0.093978 | 0.47057 | 1.0434 | -0.052961
554.0222 | 3.6361 | 2.4069 | 1.1177 | 0.012216 | 0.073428 | 0.14313 | -0.002711 | 82.8036 | 0.1003 | 54.2709 | 554.0222 | 23.7594 | 15.7271 | 7.3036 | 0.079823 | 0.4798 | 0.93522 | -0.017714
554.6913 | 3.7505 | 2.5074 | 1.175 | 0.013145 | 0.075276 | 0.1591 | -0.002238 | 85.1823 | 0.10121 | 54.5388 | 554.6913 | 23.7267 | 15.8622 | 7.4333 | 0.083158 | 0.47621 | 1.0065 | -0.014158
555.3605 | 3.8802 | 2.6072 | 1.2376 | 0.012017 | 0.082146 | 0.1523 | -0.001314 | 87.8474 | 0.10023 | 54.8068 | 555.3605 | 24.1516 | 16.2279 | 7.7035 | 0.074797 | 0.5113 | 0.94799 | -0.0081787
556.0297 | 4.0247 | 2.7134 | 1.3042 | 0.014636 | 0.078992 | 0.16184 | -0.00128 | 90.2165 | 0.10069 | 55.075 | 556.0297 | 24.4013 | 16.451 | 7.907 | 0.088737 | 0.47892 | 0.98123 | -0.0077606
556.6989 | 4.1666 | 2.8271 | 1.373 | 0.010605 | 0.083284 | 0.17636 | -0.002372 | 92.2992 | 0.10007 | 55.3432 | 556.6989 | 24.9664 | 16.9401 | 8.2269 | 0.063546 | 0.49904 | 1.0568 | -0.014213
557.368 | 4.3209 | 2.9389 | 1.443 | 0.01132 | 0.081776 | 0.16643 | -0.001172 | 93.9503 | 0.099844 | 55.6117 | 557.368 | 25.6164 | 17.4236 | 8.5548 | 0.067111 | 0.48481 | 0.98671 | -0.0069482
558.0372 | 4.474 | 3.0503 | 1.5146 | 0.011594 | 0.082375 | 0.16545 | 0.006156 | 95.5645 | 0.099446 | 55.8804 | 558.0372 | 26.3069 | 17.936 | 8.9058 | 0.068172 | 0.48436 | 0.97287 | 0.036197
558.7064 | 4.6337 | 3.172 | 1.5927 | 0.01196 | 0.080747 | 0.16778 | -0.000123 | 96.9491 | 0.098241 | 56.1493 | 558.7064 | 27.3169 | 18.7 | 9.3895 | 0.070508 | 0.47603 | 0.98911 | -0.00072512
559.3755 | 4.7919 | 3.295 | 1.6722 | 0.012214 | 0.084793 | 0.16265 | -0.002329 | 98.6477 | 0.1008 | 56.4185 | 559.3755 | 27.1873 | 18.6943 | 9.4874 | 0.069297 | 0.48108 | 0.9228 | -0.013214
560.0447 | 4.9762 | 3.4256 | 1.7588 | 0.013396 | 0.086845 | 0.17331 | -0.000913 | 100.5459 | 0.10085 | 56.688 | 560.0447 | 27.8186 | 19.1498 | 9.8323 | 0.074887 | 0.48549 | 0.96885 | -0.0051039
560.7139 | 5.1431 | 3.5717 | 1.8428 | 0.013856 | 0.09312 | 0.18222 | 0.00095 | 102.5997 | 0.099688 | 56.9579 | 560.7139 | 28.6412 | 19.8901 | 10.2624 | 0.077162 | 0.51857 | 1.0147 | 0.0052904
561.3831 | 5.335 | 3.7058 | 1.9349 | 0.012591 | 0.090853 | 0.18141 | -0.00126 | 104.819 | 0.098949 | 57.2281 | 561.3831 | 29.4373 | 20.4476 | 10.6763 | 0.069474 | 0.5013 | 1.001 | -0.0069523
562.0523 | 5.5185 | 3.8508 | 2.0298 | 0.014532 | 0.092386 | 0.1709 | -0.000303 | 107.0282 | 0.099425 | 57.4986 | 562.0523 | 29.8185 | 20.8069 | 10.9676 | 0.078521 | 0.49919 | 0.92341 | -0.0016372
562.7214 | 5.7095 | 4.005 | 2.1295 | 0.012022 | 0.094637 | 0.1892 | -0.003427 | 109.4478 | 0.099352 | 57.7694 | 562.7214 | 30.3329 | 21.2776 | 11.3134 | 0.063869 | 0.50278 | 1.0051 | -0.018207
563.3906 | 5.9083 | 4.1595 | 2.2294 | 0.013027 | 0.10115 | 0.20233 | -0.003174 | 111.8687 | 0.099626 | 58.0406 | 563.3906 | 30.7689 | 21.6618 | 11.6101 | 0.067841 | 0.52677 | 1.0537 | -0.016529
564.0598 | 6.1196 | 4.3128 | 2.3351 | 0.011752 | 0.093314 | 0.18195 | -0.004963 | 114.3769 | 0.10073 | 58.3119 | 564.0598 | 30.9731 | 21.8284 | 11.8185 | 0.059481 | 0.47229 | 0.9209 | -0.025119
564.729 | 6.3392 | 4.4838 | 2.4399 | 0.011177 | 0.10059 | 0.20741 | 0.00232 | 116.6092 | 0.099281 | 58.5836 | 564.729 | 32.0782 | 22.6895 | 12.3466 | 0.056559 | 0.509 | 1.0496 | 0.01174
565.3982 | 6.5466 | 4.6526 | 2.5522 | 0.012557 | 0.10346 | 0.19654 | 0.004821 | 119.3168 | 0.099457 | 58.8554 | 565.3982 | 32.4684 | 23.0748 | 12.6579 | 0.062278 | 0.51312 | 0.97476 | 0.02391
566.0673 | 6.7633 | 4.8198 | 2.6639 | 0.016891 | 0.10218 | 0.20827 | 0.005278 | 121.7778 | 0.099543 | 59.1275 | 566.0673 | 32.9889 | 23.5093 | 12.9938 | 0.082389 | 0.49838 | 1.0159 | 0.025744
566.7365 | 6.9793 | 5.0112 | 2.7848 | 0.01537 | 0.11102 | 0.20799 | -0.000612 | 124.5166 | 0.099297 | 59.3998 | 566.7365 | 33.5298 | 24.0749 | 13.3788 | 0.073841 | 0.53336 | 0.99922 | -0.0029402
567.4057 | 7.2128 | 5.1848 | 2.9074 | 0.012816 | 0.12068 | 0.21159 | 0.000767 | 127.1645 | 0.099299 | 59.6723 | 567.4057 | 34.0852 | 24.5015 | 13.7392 | 0.060564 | 0.57027 | 0.99992 | 0.0036246
568.0749 | 7.4517 | 5.3758 | 3.0384 | 0.015227 | 0.11799 | 0.22192 | -0.003283 | 130.0725 | 0.098519 | 59.9449 | 568.0749 | 34.858 | 25.147 | 14.2134 | 0.071229 | 0.55194 | 1.0381 | -0.015357
568.7441 | 7.6998 | 5.5685 | 3.1728 | 0.014182 | 0.12336 | 0.22874 | 0.002433 | 133.0278 | 0.098017 | 60.2177 | 568.7441 | 35.5599 | 25.7171 | 14.653 | 0.065497 | 0.5697 | 1.0564 | 0.011236
569.4133 | 7.9563 | 5.7655 | 3.306 | 0.015602 | 0.12884 | 0.23319 | -0.002736 | 136.1532 | 0.098147 | 60.4906 | 569.4133 | 36.0161 | 26.099 | 14.9652 | 0.070626 | 0.58323 | 1.0556 | -0.012385
570.0824 | 8.2108 | 5.9643 | 3.4453 | 0.013143 | 0.13732 | 0.23355 | 0.007966 | 138.9828 | 0.098514 | 60.7636 | 570.0824 | 36.4392 | 26.4695 | 15.2899 | 0.058328 | 0.60944 | 1.0365 | 0.035353
570.7516 | 8.4635 | 6.1671 | 3.5868 | 0.01543 | 0.14393 | 0.23325 | -0.001754 | 141.8379 | 0.098538 | 61.0368 | 570.7516 | 36.9611 | 26.9328 | 15.664 | 0.067385 | 0.62854 | 1.0186 | -0.00766
571.4208 | 8.7225 | 6.3781 | 3.7317 | 0.016521 | 0.15394 | 0.25097 | 0.000289 | 144.7964 | 0.098325 | 61.31 | 571.4208 | 37.562 | 27.4661 | 16.0698 | 0.071145 | 0.66293 | 1.0808 | 0.0012445
572.09 | 8.9878 | 6.5921 | 3.8777 | 0.016348 | 0.16745 | 0.25891 | 0.003878 | 147.5887 | 0.097938 | 61.5832 | 572.09 | 38.2925 | 28.0856 | 16.5208 | 0.06965 | 0.71343 | 1.1031 | 0.016522
572.7592 | 9.2343 | 6.7954 | 4.0305 | 0.014706 | 0.17856 | 0.25364 | -0.004581 | 150.4083 | 0.098068 | 61.8564 | 572.7592 | 38.725 | 28.4971 | 16.9025 | 0.061671 | 0.74882 | 1.0637 | -0.019211
573.4283 | 9.5082 | 7.0195 | 4.1876 | 0.01382 | 0.18795 | 0.2743 | -0.001023 | 153.1156 | 0.097746 | 62.1295 | 573.4283 | 39.471 | 29.1398 | 17.384 | 0.057371 | 0.78023 | 1.1387 | -0.0042467
574.0975 | 9.7718 | 7.2427 | 4.348 | 0.01276 | 0.20241 | 0.25679 | -0.00447 | 156.0241 | 0.097558 | 62.4025 | 574.0975 | 40.0611 | 29.6925 | 17.8252 | 0.052312 | 0.82983 | 1.0527 | -0.018326
574.7667 | 10.0753 | 7.4578 | 4.5228 | 0.015348 | 0.22129 | 0.27254 | 0.004579 | 159.1411 | 0.097993 | 62.6754 | 574.7667 | 40.4926 | 29.9732 | 18.1771 | 0.061684 | 0.88938 | 1.0953 | 0.018403
575.4359 | 10.359 | 7.6966 | 4.6956 | 0.017703 | 0.23948 | 0.30221 | -0.002911 | 162.1796 | 0.098021 | 62.9481 | 575.4359 | 41.0189 | 30.4765 | 18.5935 | 0.070099 | 0.94827 | 1.1967 | -0.011527
576.105 | 10.6594 | 7.9361 | 4.8766 | 0.015552 | 0.26781 | 0.29536 | -0.000648 | 165.3748 | 0.097966 | 63.2206 | 576.105 | 41.5954 | 30.9685 | 19.0298 | 0.060688 | 1.0451 | 1.1526 | -0.0025287
576.7742 | 10.98 | 8.1762 | 5.0547 | 0.019123 | 0.30238 | 0.30586 | -0.00519 | 168.8502 | 0.096948 | 63.4928 | 576.7742 | 42.5881 | 31.7129 | 19.6056 | 0.074172 | 1.1729 | 1.1863 | -0.02013
577.4434 | 11.2882 | 8.4145 | 5.2388 | 0.016184 | 0.34044 | 0.30139 | -0.003615 | 171.8503 | 0.097391 | 63.7647 | 577.4434 | 43.0066 | 32.0581 | 19.9591 | 0.061659 | 1.297 | 1.1483 | -0.013773
578.1125 | 11.62 | 8.6695 | 5.4305 | 0.018069 | 0.38719 | 0.31141 | -0.006368 | 175.2519 | 0.097924 | 64.0363 | 578.1125 | 43.3591 | 32.3493 | 20.2635 | 0.067423 | 1.4448 | 1.162 | -0.023762
578.7817 | 11.9181 | 8.9235 | 5.6204 | 0.018391 | 0.44026 | 0.33705 | -3.6e-05 | 178.4024 | 0.097733 | 64.3074 | 578.7817 | 43.957 | 32.912 | 20.7293 | 0.067831 | 1.6238 | 1.2431 | -0.00013278
579.4509 | 12.2391 | 9.178 | 5.8113 | 0.018057 | 0.49967 | 0.34429 | -0.003211 | 181.6712 | 0.098166 | 64.5782 | 579.4509 | 44.319 | 33.2344 | 21.0433 | 0.065386 | 1.8094 | 1.2467 | -0.011627
580.1201 | 12.5651 | 9.4439 | 6.0075 | 0.018952 | 0.57686 | 0.35689 | -0.002994 | 184.9611 | 0.097444 | 64.8485 | 580.1201 | 45.2095 | 33.9794 | 21.615 | 0.06819 | 2.0756 | 1.2841 | -0.010772
580.7892 | 12.8886 | 9.7004 | 6.1992 | 0.018106 | 0.66743 | 0.36639 | 0.002645 | 188.0761 | 0.097541 | 65.1183 | 580.7892 | 45.7497 | 34.4328 | 22.0049 | 0.06427 | 2.3691 | 1.3006 | 0.0093888
581.4584 | 13.2108 | 9.9624 | 6.3998 | 0.01756 | 0.7705 | 0.37104 | 0.006661 | 191.0078 | 0.09765 | 65.3877 | 581.4584 | 46.3131 | 34.9251 | 22.4357 | 0.06156 | 2.7011 | 1.3008 | 0.023351
582.1276 | 13.5148 | 10.2347 | 6.6064 | 0.019751 | 0.90166 | 0.39291 | -0.003871 | 194.3529 | 0.097341 | 65.6569 | 582.1276 | 46.9036 | 35.5197 | 22.9277 | 0.068546 | 3.1292 | 1.3636 | -0.013434
582.7968 | 13.8584 | 10.5052 | 6.81 | 0.019744 | 1.0582 | 0.38606 | -0.006048 | 197.4448 | 0.097267 | 65.9259 | 582.7968 | 47.5728 | 36.0623 | 23.3773 | 0.067777 | 3.6327 | 1.3253 | -0.020762
583.466 | 14.1834 | 10.7808 | 7.0183 | 0.01887 | 1.2249 | 0.41059 | -0.010273 | 200.7858 | 0.096955 | 66.1949 | 583.466 | 48.228 | 36.6583 | 23.8644 | 0.064164 | 4.1651 | 1.3962 | -0.034931
584.1352 | 14.5548 | 11.0791 | 7.2407 | 0.018524 | 1.434 | 0.43303 | -0.004525 | 204.118 | 0.097366 | 66.4639 | 584.1352 | 48.675 | 37.0513 | 24.2145 | 0.061949 | 4.7956 | 1.4482 | -0.015133
584.8043 | 14.8858 | 11.3583 | 7.4627 | 0.01982 | 1.6874 | 0.44939 | -0.002171 | 207.2265 | 0.09691 | 66.7331 | 584.8043 | 49.4653 | 37.7433 | 24.7985 | 0.065861 | 5.6072 | 1.4933 | -0.0072142
585.4735 | 15.2728 | 11.6696 | 7.6994 | 0.018878 | 1.976 | 0.46415 | -0.002997 | 210.9148 | 0.096832 | 67.0025 | 585.4735 | 50.1052 | 38.2842 | 25.2594 | 0.061933 | 6.4828 | 1.5227 | -0.0098322
586.1427 | 15.6467 | 11.9854 | 7.9427 | 0.021412 | 2.3136 | 0.48015 | -0.00063 | 214.2682 | 0.097268 | 67.2723 | 586.1427 | 50.5044 | 38.6863 | 25.6375 | 0.069114 | 7.4677 | 1.5498 | -0.0020335
586.8119 | 16.0342 | 12.314 | 8.1827 | 0.021289 | 2.72 | 0.50924 | 0.002457 | 218.0621 | 0.097519 | 67.5426 | 586.8119 | 50.928 | 39.112 | 25.9899 | 0.067618 | 8.6394 | 1.6174 | 0.0078039
587.4811 | 16.4229 | 12.6566 | 8.4515 | 0.020477 | 3.2041 | 0.51909 | -0.000238 | 222.1499 | 0.096786 | 67.8134 | 587.4811 | 51.797 | 39.9183 | 26.6555 | 0.064584 | 10.1057 | 1.6372 | -0.00075064
588.1502 | 16.8626 | 13.0167 | 8.7219 | 0.020095 | 3.7647 | 0.56068 | -0.003485 | 226.3428 | 0.097203 | 68.0849 | 588.1502 | 52.1829 | 40.2812 | 26.9907 | 0.062186 | 11.6503 | 1.7351 | -0.010785
588.8194 | 17.2953 | 13.39 | 9.0074 | 0.022253 | 4.4326 | 0.59657 | -0.003005 | 231.1803 | 0.096911 | 68.3573 | 588.8194 | 52.7702 | 40.8547 | 27.4828 | 0.067897 | 13.5243 | 1.8202 | -0.0091686
589.4886 | 17.7806 | 13.7914 | 9.3113 | 0.021361 | 5.2262 | 0.60689 | 0.001591 | 236.1353 | 0.097507 | 68.6305 | 589.4886 | 52.999 | 41.1084 | 27.7542 | 0.063671 | 15.5779 | 1.809 | 0.0047423
590.1578 | 18.2668 | 14.1967 | 9.6209 | 0.02086 | 6.1511 | 0.6165 | 0.003651 | 241.2243 | 0.097163 | 68.9048 | 590.1578 | 53.7016 | 41.7363 | 28.2839 | 0.061325 | 18.0833 | 1.8124 | 0.010733
590.827 | 18.7668 | 14.6216 | 9.9384 | 0.023204 | 7.2065 | 0.682 | -0.005958 | 246.5476 | 0.09705 | 69.18 | 590.827 | 54.2594 | 42.2746 | 28.7342 | 0.067088 | 20.8356 | 1.9718 | -0.017226
591.4961 | 19.2807 | 15.0628 | 10.2662 | 0.025109 | 8.441 | 0.70585 | -0.005242 | 251.6823 | 0.09723 | 69.4561 | 591.4961 | 54.7243 | 42.7525 | 29.1383 | 0.071267 | 23.9579 | 2.0034 | -0.014878
592.1653 | 19.7945 | 15.5121 | 10.5981 | 0.024133 | 9.878 | 0.75429 | 0.004701 | 257.1567 | 0.096728 | 69.733 | 592.1653 | 55.4922 | 43.4869 | 29.7109 | 0.067655 | 27.6922 | 2.1146 | 0.013179
592.8345 | 20.3277 | 15.974 | 10.9352 | 0.022708 | 11.5256 | 0.78078 | -0.000797 | 262.4586 | 0.096856 | 70.0104 | 592.8345 | 55.9839 | 43.9936 | 30.1163 | 0.062539 | 31.7424 | 2.1503 | -0.002195
593.5037 | 20.8644 | 16.4215 | 11.2829 | 0.023389 | 13.4183 | 0.85044 | -0.000418 | 268.0214 | 0.096686 | 70.2882 | 593.5037 | 56.5919 | 44.5412 | 30.6032 | 0.06344 | 36.3952 | 2.3067 | -0.0011338
594.1729 | 21.4206 | 16.8931 | 11.6253 | 0.026618 | 15.5547 | 0.89855 | -0.001946 | 273.5027 | 0.096903 | 70.5661 | 594.1729 | 57.0336 | 44.9789 | 30.953 | 0.070872 | 41.4153 | 2.3925 | -0.0051813
594.842 | 21.9407 | 17.3478 | 11.9643 | 0.026003 | 17.9539 | 0.96071 | 0.006267 | 278.6468 | 0.096418 | 70.8442 | 594.842 | 57.8552 | 45.7444 | 31.5485 | 0.068567 | 47.3425 | 2.5333 | 0.016525
595.5112 | 22.4647 | 17.7961 | 12.3086 | 0.026411 | 20.6136 | 1.0181 | -0.004151 | 283.5783 | 0.096636 | 71.122 | 595.5112 | 58.3033 | 46.1867 | 31.9447 | 0.068545 | 53.4991 | 2.6423 | -0.010773
596.1804 | 22.9857 | 18.24 | 12.6526 | 0.025931 | 23.6136 | 1.1091 | -0.001559 | 288.4799 | 0.096635 | 71.3996 | 596.1804 | 58.871 | 46.7164 | 32.406 | 0.066415 | 60.4793 | 2.8406 | -0.0039929
596.8495 | 23.5182 | 18.6771 | 12.9887 | 0.027726 | 26.992 | 1.2012 | 0.001841 | 293.1394 | 0.096796 | 71.6767 | 596.8495 | 59.4087 | 47.1797 | 32.8106 | 0.070038 | 68.1838 | 3.0344 | 0.0046505
597.5187 | 24.0264 | 19.1317 | 13.3423 | 0.02845 | 30.6916 | 1.2933 | -0.001561 | 297.9874 | 0.096207 | 71.9532 | 597.5187 | 60.3023 | 48.0175 | 33.487 | 0.071405 | 77.0308 | 3.2459 | -0.0039178
598.1879 | 24.5474 | 19.5754 | 13.6896 | 0.028436 | 34.8429 | 1.3823 | 0.000592 | 303.15 | 0.096482 | 72.2289 | 598.1879 | 60.6193 | 48.3409 | 33.8061 | 0.070222 | 86.0439 | 3.4135 | 0.0014619
598.8571 | 25.0954 | 20.0481 | 14.0472 | 0.028681 | 39.5129 | 1.51 | 0.003353 | 307.8508 | 0.096491 | 72.5036 | 598.8571 | 61.2525 | 48.9332 | 34.2862 | 0.070004 | 96.4426 | 3.6856 | 0.008184
599.5262 | 25.621 | 20.4941 | 14.3924 | 0.028794 | 44.5852 | 1.6409 | 0.006382 | 312.5724 | 0.096273 | 72.7771 | 599.5262 | 61.9638 | 49.5645 | 34.8076 | 0.069638 | 107.8283 | 3.9684 | 0.015435
600.1954 | 26.179 | 20.9729 | 14.7702 | 0.028973 | 50.0668 | 1.8125 | 0.004834 | 317.3765 | 0.096288 | 73.0494 | 600.1954 | 62.5777 | 50.1334 | 35.3064 | 0.069257 | 119.6788 | 4.3326 | 0.011555
600.8646 | 26.7127 | 21.4342 | 15.1165 | 0.029485 | 55.9086 | 1.9728 | 0.000445 | 322.1934 | 0.096233 | 73.3202 | 600.8646 | 63.1686 | 50.6863 | 35.7467 | 0.069724 | 132.2094 | 4.6652 | 0.0010523
601.5338 | 27.2639 | 21.9107 | 15.4873 | 0.031646 | 62.2176 | 2.1792 | 0.001492 | 326.9264 | 0.096189 | 73.5898 | 601.5338 | 63.8017 | 51.2745 | 36.2427 | 0.074056 | 145.5987 | 5.0997 | 0.0034915
602.203 | 27.8008 | 22.379 | 15.85 | 0.030381 | 68.9213 | 2.4011 | -0.01042 | 331.9541 | 0.096247 | 73.8583 | 602.203 | 64.2676 | 51.7338 | 36.6408 | 0.070232 | 159.3265 | 5.5507 | -0.024088
602.8721 | 28.3935 | 22.8754 | 16.2284 | 0.033584 | 76.1038 | 2.6589 | -0.004311 | 336.9338 | 0.096253 | 74.1257 | 602.8721 | 64.8975 | 52.2851 | 37.0925 | 0.076761 | 173.9466 | 6.0772 | -0.0098534
603.5413 | 28.9435 | 23.3643 | 16.6046 | 0.033146 | 83.7342 | 2.9666 | 0.00103 | 341.7607 | 0.096093 | 74.3921 | 603.5413 | 65.564 | 52.9257 | 37.6135 | 0.075084 | 189.678 | 6.72 | 0.0023332
604.2105 | 29.5478 | 23.8761 | 16.9924 | 0.031301 | 91.8632 | 3.2886 | -0.001027 | 347.1956 | 0.096016 | 74.6578 | 604.2105 | 66.1732 | 53.4712 | 38.055 | 0.070099 | 205.7302 | 7.3648 | -0.0023
604.8797 | 30.1223 | 24.3835 | 17.3861 | 0.034562 | 100.4961 | 3.6586 | 0.004413 | 352.0983 | 0.095807 | 74.9228 | 604.8797 | 66.9023 | 54.1563 | 38.6149 | 0.076763 | 223.2039 | 8.1258 | 0.0098014
605.5489 | 30.723 | 24.8917 | 17.7876 | 0.032791 | 109.7204 | 4.071 | 0.003482 | 356.9936 | 0.095646 | 75.1872 | 605.5489 | 67.6517 | 54.8113 | 39.1682 | 0.072205 | 241.6031 | 8.9644 | 0.0076673
606.218 | 31.3424 | 25.4061 | 18.1894 | 0.036255 | 119.3993 | 4.5685 | 0.006142 | 361.996 | 0.095487 | 75.4511 | 606.218 | 68.4151 | 55.4571 | 39.7044 | 0.079138 | 260.6281 | 9.9723 | 0.013407
606.8872 | 31.9899 | 25.945 | 18.589 | 0.040277 | 129.6411 | 5.1066 | -0.001662 | 366.9667 | 0.09569 | 75.7148 | 606.8872 | 68.9764 | 55.9424 | 40.0815 | 0.086845 | 279.531 | 11.0108 | -0.0035836
607.5564 | 32.559 | 26.4569 | 19.0051 | 0.036761 | 140.3178 | 5.7201 | -0.007706 | 372.1982 | 0.095725 | 75.9782 | 607.5564 | 69.4318 | 56.4192 | 40.5283 | 0.078393 | 299.2269 | 12.1981 | -0.016433
608.2256 | 33.2081 | 26.9821 | 19.4058 | 0.038014 | 151.5505 | 6.4579 | 0.002079 | 377.0809 | 0.095315 | 76.2414 | 608.2256 | 70.4435 | 57.2365 | 41.1651 | 0.080638 | 321.4805 | 13.6991 | 0.0044101
608.8948 | 33.8308 | 27.5253 | 19.8146 | 0.038374 | 163.4041 | 7.2191 | -0.000647 | 382.2357 | 0.09527 | 76.5048 | 608.8948 | 71.0747 | 57.8276 | 41.6283 | 0.080619 | 343.2938 | 15.1666 | -0.0013593
609.564 | 34.4749 | 28.079 | 20.2379 | 0.036597 | 175.793 | 8.1593 | -0.000896 | 387.5237 | 0.094769 | 76.7682 | 609.564 | 72.0645 | 58.6949 | 42.3043 | 0.0765 | 367.4684 | 17.0558 | -0.001873
610.2331 | 35.1288 | 28.6251 | 20.6657 | 0.041572 | 188.809 | 9.1766 | 0.004067 | 392.7446 | 0.095069 | 77.0319 | 610.2331 | 72.4743 | 59.0565 | 42.6353 | 0.085767 | 389.5318 | 18.9323 | 0.0083906
610.9023 | 35.7944 | 29.1965 | 21.0942 | 0.040266 | 202.5759 | 10.354 | -0.001841 | 397.7454 | 0.094838 | 77.2959 | 610.9023 | 73.3475 | 59.8274 | 43.2247 | 0.08251 | 415.1047 | 21.2168 | -0.0037725
611.5715 | 36.4555 | 29.7709 | 21.5143 | 0.044195 | 217.095 | 11.6424 | 0.001103 | 403.0847 | 0.094805 | 77.5602 | 611.5715 | 73.9904 | 60.4234 | 43.6656 | 0.089699 | 440.6184 | 23.6296 | 0.0022387
612.2407 | 37.1482 | 30.3543 | 21.9572 | 0.04296 | 232.1559 | 13.158 | -1.6e-05 | 408.7007 | 0.094877 | 77.8248 | 612.2407 | 74.5575 | 60.922 | 44.0687 | 0.086222 | 465.944 | 26.4085 | -3.2112e-05
612.9099 | 37.7889 | 30.9237 | 22.4029 | 0.042941 | 247.5578 | 14.8412 | 0.000453 | 413.9326 | 0.09457 | 78.0897 | 612.9099 | 75.3836 | 61.6886 | 44.6907 | 0.085661 | 493.8438 | 29.6062 | 0.00090367
613.579 | 38.456 | 31.5088 | 22.8353 | 0.043205 | 262.9005 | 16.7021 | 0.005021 | 419.1693 | 0.094455 | 78.3549 | 613.579 | 76.1052 | 62.3565 | 45.1915 | 0.085504 | 520.2855 | 33.0538 | 0.0099367
614.2482 | 39.1232 | 32.0744 | 23.2703 | 0.045262 | 277.9155 | 18.864 | 0.005276 | 424.2376 | 0.094226 | 78.6203 | 614.2482 | 76.947 | 63.0834 | 45.7677 | 0.089021 | 546.5998 | 37.1015 | 0.010377
614.9174 | 39.8165 | 32.6767 | 23.7297 | 0.046802 | 292.5164 | 21.2182 | -0.007617 | 429.7109 | 0.094472 | 78.886 | 614.9174 | 77.372 | 63.4978 | 46.1118 | 0.090946 | 568.4218 | 41.2315 | -0.014801
615.5865 | 40.4676 | 33.2707 | 24.1762 | 0.04874 | 306.0187 | 23.8808 | 0.005564 | 434.7552 | 0.094477 | 79.1518 | 615.5865 | 77.9823 | 64.1138 | 46.5884 | 0.093923 | 589.7073 | 46.0191 | 0.010722
616.2557 | 41.1385 | 33.8773 | 24.6316 | 0.047181 | 318.5345 | 26.9588 | -0.002588 | 440.203 | 0.094351 | 79.4179 | 616.2557 | 78.6627 | 64.7783 | 47.0991 | 0.090217 | 609.0832 | 51.5491 | -0.0049486
616.9249 | 41.818 | 34.4829 | 25.0774 | 0.050083 | 329.7235 | 30.3482 | -0.004634 | 445.6701 | 0.094079 | 79.6842 | 616.9249 | 79.475 | 65.5345 | 47.6594 | 0.095183 | 626.6383 | 57.6766 | -0.0088069
617.5941 | 42.5524 | 35.1162 | 25.5533 | 0.049607 | 339.4816 | 34.086 | -0.005576 | 451.0542 | 0.094017 | 79.9507 | 617.5941 | 80.2252 | 66.2056 | 48.1763 | 0.093525 | 640.0344 | 64.2633 | -0.010513
618.2632 | 43.2656 | 35.7074 | 26.0174 | 0.054039 | 347.2973 | 38.3268 | -0.003822 | 456.5811 | 0.094188 | 80.2173 | 618.2632 | 80.7042 | 66.6056 | 48.5307 | 0.1008 | 647.8205 | 71.4918 | -0.0071293
618.9324 | 43.9876 | 36.358 | 26.482 | 0.053751 | 353.2745 | 42.9638 | -0.010007 | 461.9811 | 0.094124 | 80.4841 | 618.9324 | 81.4172 | 67.2955 | 49.0158 | 0.099488 | 653.8805 | 79.5222 | -0.018522
619.6016 | 44.6737 | 36.971 | 26.9632 | 0.054716 | 356.7109 | 48.1985 | 0.003335 | 467.603 | 0.094053 | 80.751 | 619.6016 | 82.0258 | 67.8828 | 49.5073 | 0.10046 | 654.9596 | 88.4977 | 0.0061234
620.2708 | 45.3831 | 37.6447 | 27.4299 | 0.059356 | 357.6254 | 54.0079 | -0.003228 | 472.7934 | 0.093867 | 81.0181 | 620.2708 | 82.8501 | 68.7232 | 50.0753 | 0.10836 | 652.8713 | 98.5953 | -0.0058929
620.9399 | 46.1187 | 38.2303 | 27.8879 | 0.05818 | 355.6793 | 60.3818 | 0.002275 | 478.1223 | 0.093969 | 81.2852 | 620.9399 | 83.4379 | 69.1662 | 50.4547 | 0.10526 | 643.4942 | 109.2427 | 0.0041159
621.6091 | 46.8087 | 38.8606 | 28.3365 | 0.05992 | 350.9461 | 67.4542 | -0.002181 | 483.5041 | 0.093681 | 81.5524 | 621.6091 | 84.2771 | 69.9669 | 51.0187 | 0.10788 | 631.8637 | 121.4484 | -0.0039268
622.2783 | 47.4966 | 39.4731 | 28.8044 | 0.061188 | 343.7855 | 75.2998 | 0.00332 | 488.7236 | 0.093847 | 81.8197 | 622.2783 | 84.7299 | 70.4166 | 51.3845 | 0.10915 | 613.2839 | 134.3284 | 0.0059226
622.9475 | 48.2297 | 40.1044 | 29.2611 | 0.060227 | 334.5235 | 83.8136 | 0.00539 | 493.9561 | 0.093664 | 82.0869 | 622.9475 | 85.5709 | 71.1548 | 51.9162 | 0.10686 | 593.5247 | 148.7054 | 0.0095632
623.6167 | 48.8652 | 40.6939 | 29.6726 | 0.065084 | 323.2685 | 92.8526 | 0.004013 | 498.545 | 0.093761 | 82.354 | 623.6167 | 86.0907 | 71.6945 | 52.2773 | 0.11467 | 569.5346 | 163.5878 | 0.0070701
624.2859 | 49.6268 | 41.3226 | 30.1148 | 0.064052 | 311.2221 | 103.2274 | 0.006489 | 504.1999 | 0.093636 | 82.621 | 624.2859 | 86.8487 | 72.316 | 52.702 | 0.11209 | 544.6496 | 180.6516 | 0.011356
624.955 | 50.28 | 41.9486 | 30.5466 | 0.065852 | 298.4468 | 114.3182 | -0.000773 | 509.2252 | 0.093602 | 82.8878 | 624.955 | 87.4359 | 72.9477 | 53.1199 | 0.11452 | 518.9926 | 198.7969 | -0.0013442
625.6242 | 51.0008 | 42.5831 | 30.9941 | 0.070443 | 285.3903 | 126.8415 | 0.0031 | 514.2351 | 0.093292 | 83.1545 | 625.6242 | 88.401 | 73.8104 | 53.723 | 0.1221 | 494.6746 | 219.8577 | 0.0053733
626.2934 | 51.708 | 43.264 | 31.4541 | 0.070801 | 272.235 | 140.3495 | 0.001556 | 519.7594 | 0.093291 | 83.4209 | 626.2934 | 88.9588 | 74.4316 | 54.1139 | 0.12181 | 468.3552 | 241.4583 | 0.002677
626.9626 | 52.4621 | 43.8716 | 31.8879 | 0.072511 | 259.0637 | 155.3213 | 0.001155 | 525.0607 | 0.093298 | 83.687 | 626.9626 | 89.6234 | 74.9477 | 54.4755 | 0.12387 | 442.5698 | 265.3422 | 0.0019731
627.6318 | 53.2089 | 44.562 | 32.3438 | 0.071502 | 246.0009 | 171.9208 | -0.00085 | 530.5063 | 0.093328 | 83.9528 | 627.6318 | 90.2232 | 75.5612 | 54.8435 | 0.12124 | 417.1294 | 291.5161 | -0.0014413
628.3009 | 53.9354 | 45.2007 | 32.7693 | 0.074911 | 232.9794 | 189.9656 | 8.1e-05 | 535.6681 | 0.093315 | 84.2181 | 628.3009 | 90.873 | 76.1562 | 55.2112 | 0.12621 | 392.5346 | 320.0629 | 0.00013647
628.9701 | 54.6887 | 45.862 | 33.2296 | 0.07537 | 220.3186 | 209.261 | -0.000823 | 540.9779 | 0.093134 | 84.4831 | 628.9701 | 91.7024 | 76.9016 | 55.7196 | 0.12638 | 369.4313 | 350.89 | -0.00138
629.6393 | 55.4526 | 46.5586 | 33.6911 | 0.077776 | 208.1351 | 230.5542 | -0.001294 | 546.7879 | 0.093194 | 84.7476 | 629.6393 | 92.2239 | 77.4322 | 56.0321 | 0.12935 | 346.1521 | 383.4375 | -0.0021521
630.3085 | 56.2056 | 47.2579 | 34.141 | 0.079955 | 196.4779 | 253.4014 | 0.006155 | 552.439 | 0.093162 | 85.0116 | 630.3085 | 92.8396 | 78.06 | 56.3935 | 0.13207 | 324.5393 | 418.5647 | 0.010167
630.9777 | 57.0354 | 47.9171 | 34.6039 | 0.081816 | 185.2629 | 278.2956 | -0.007026 | 557.8683 | 0.092945 | 85.2751 | 630.9777 | 93.8011 | 78.805 | 56.91 | 0.13456 | 304.6857 | 457.6886 | -0.011555
631.6469 | 57.7978 | 48.6418 | 35.0684 | 0.085527 | 174.6838 | 305.0789 | -0.002961 | 563.8264 | 0.092884 | 85.5381 | 631.6469 | 94.4028 | 79.4481 | 57.2783 | 0.13969 | 285.3164 | 498.2947 | -0.0048363
632.316 | 58.5777 | 49.3793 | 35.5398 | 0.092749 | 164.8457 | 334.8841 | -0.002966 | 569.8357 | 0.092838 | 85.8005 | 632.316 | 95.005 | 80.0863 | 57.6407 | 0.15043 | 267.3569 | 543.1356 | -0.0048104
632.975 | 59.4272 | 50.0858 | 36.0123 | 0.093318 | 155.5002 | 366.5039 | -0.00297 | 575.5849 | 0.092731 | 86.0585 | 632.975 | 95.8171 | 80.7555 | 58.0642 | 0.15046 | 250.7199 | 590.9302 | -0.0047887
633.6341 | 60.2108 | 50.8085 | 36.4578 | 0.099055 | 146.8242 | 400.5365 | -0.000941 | 581.3036 | 0.092567 | 86.316 | 633.6341 | 96.5846 | 81.5023 | 58.4823 | 0.15889 | 235.5217 | 642.5033 | -0.0015095
634.2932 | 61.0618 | 51.5327 | 36.9133 | 0.10536 | 138.6425 | 437.3501 | -0.002776 | 587.5014 | 0.092557 | 86.573 | 634.2932 | 97.2149 | 82.0439 | 58.7687 | 0.16773 | 220.7292 | 696.2939 | -0.0044196
634.9522 | 61.866 | 52.2723 | 37.3589 | 0.11761 | 130.9652 | 476.5059 | -0.011184 | 593.0507 | 0.092592 | 86.8296 | 634.9522 | 97.8262 | 82.6561 | 59.0741 | 0.18597 | 207.0901 | 753.4801 | -0.017685
635.6113 | 62.6857 | 52.9963 | 37.8004 | 0.13519 | 123.746 | 518.6837 | -0.005985 | 599.0012 | 0.092644 | 87.0857 | 635.6113 | 98.3715 | 83.1662 | 59.3196 | 0.21215 | 194.1924 | 813.9611 | -0.0093922
636.2703 | 63.4927 | 53.7434 | 38.2218 | 0.16102 | 116.9552 | 563.4438 | -0.003507 | 605.0582 | 0.092272 | 87.3413 | 636.2703 | 99.3288 | 84.0768 | 59.7946 | 0.25191 | 182.9663 | 881.4587 | -0.0054864
636.9294 | 64.3708 | 54.4991 | 38.6576 | 0.20563 | 110.4965 | 611.3742 | -0.00058 | 611.0413 | 0.092466 | 87.5965 | 636.9294 | 99.7979 | 84.4932 | 59.9332 | 0.3188 | 171.3094 | 947.8502 | -0.00089921
637.5884 | 65.2439 | 55.2825 | 39.0983 | 0.28076 | 104.4844 | 662.5839 | -0.005586 | 617.6428 | 0.092385 | 87.8513 | 637.5884 | 100.4495 | 85.1129 | 60.1958 | 0.43225 | 160.8641 | 1020.1137 | -0.0086002
638.2474 | 66.0984 | 56.0499 | 39.5129 | 0.39438 | 98.8447 | 715.9568 | -0.005146 | 623.9773 | 0.092297 | 88.1057 | 638.2474 | 101.1206 | 85.748 | 60.4488 | 0.60334 | 151.2176 | 1095.3064 | -0.0078726
638.9065 | 66.9717 | 56.8284 | 39.9153 | 0.57156 | 93.6024 | 771.6108 | -0.001749 | 629.9234 | 0.092369 | 88.3596 | 638.9065 | 101.7019 | 86.2985 | 60.6146 | 0.86797 | 142.1427 | 1171.7529 | -0.002656
639.5656 | 67.8191 | 57.6044 | 40.2961 | 0.83 | 88.8409 | 829.8365 | 0.002562 | 636.1727 | 0.092129 | 88.6131 | 639.5656 | 102.5367 | 87.0929 | 60.9242 | 1.2549 | 134.3198 | 1254.6418 | 0.0038735
640.2246 | 68.7103 | 58.3919 | 40.6843 | 1.1856 | 84.5858 | 889.1162 | -0.001528 | 642.3697 | 0.092126 | 88.8662 | 640.2246 | 103.1796 | 87.6848 | 61.0939 | 1.7804 | 127.0191 | 1335.1503 | -0.0022945
640.8837 | 69.5442 | 59.1516 | 41.0391 | 1.6766 | 80.6003 | 950.1524 | -0.00675 | 648.6024 | 0.091958 | 89.1189 | 640.8837 | 103.911 | 88.3826 | 61.3195 | 2.5051 | 120.4307 | 1419.6907 | -0.010086
641.5427 | 70.4447 | 59.9711 | 41.4036 | 2.3171 | 77.0888 | 1011.9807 | -0.009017 | 654.3558 | 0.091919 | 89.3713 | 641.5427 | 104.6707 | 89.1085 | 61.5198 | 3.4428 | 114.5428 | 1503.6573 | -0.013398
642.2018 | 71.2924 | 60.7432 | 41.7554 | 3.1243 | 73.8306 | 1073.8086 | -0.000612 | 660.8383 | 0.092111 | 89.6233 | 642.2018 | 104.9687 | 89.4364 | 61.4794 | 4.6002 | 108.7059 | 1581.0429 | -0.00090109
642.8608 | 72.2086 | 61.5395 | 42.0543 | 4.1142 | 70.8092 | 1134.8459 | 9.7e-05 | 666.8781 | 0.091904 | 89.8751 | 642.8608 | 105.8881 | 90.2427 | 61.6693 | 6.0332 | 103.836 | 1664.1611 | 0.00014224
643.5199 | 73.0762 | 62.3386 | 42.3424 | 5.2906 | 67.9904 | 1194.3402 | 0.001485 | 673.295 | 0.091762 | 90.1265 | 643.5199 | 106.601 | 90.9373 | 61.7675 | 7.7177 | 99.182 | 1742.2607 | 0.0021663
644.179 | 73.9571 | 63.1263 | 42.63 | 6.6275 | 65.3903 | 1251.3591 | -0.006678 | 679.066 | 0.09165 | 90.3777 | 644.179 | 107.3977 | 91.6697 | 61.9056 | 9.6242 | 94.9573 | 1817.176 | -0.0096975
644.838 | 74.8432 | 63.9235 | 42.8838 | 8.1667 | 62.8616 | 1304.2926 | -0.004424 | 685.4429 | 0.091724 | 90.6287 | 644.838 | 107.8857 | 92.1451 | 61.8165 | 11.7723 | 90.6144 | 1880.124 | -0.0063771
645.497 | 75.7264 | 64.7455 | 43.1232 | 9.805 | 60.4716 | 1353.4622 | 0.001011 | 691.9874 | 0.091688 | 90.8795 | 645.497 | 108.4685 | 92.7397 | 61.7684 | 14.0445 | 86.6178 | 1938.6626 | 0.0014481
646.1561 | 76.5928 | 65.5212 | 43.3236 | 11.552 | 58.0863 | 1397.2485 | -0.004875 | 697.7689 | 0.091491 | 91.1301 | 646.1561 | 109.3347 | 93.5302 | 61.8435 | 16.4903 | 82.917 | 1994.5445 | -0.006959
646.8151 | 77.526 | 66.3515 | 43.5532 | 13.4129 | 55.9039 | 1436.7247 | 0.00332 | 704.8768 | 0.091421 | 91.3806 | 646.8151 | 109.9367 | 94.0905 | 61.7611 | 19.0204 | 79.2752 | 2037.3642 | 0.004708
647.4742 | 78.4114 | 67.1876 | 43.7431 | 15.3439 | 53.823 | 1469.6177 | -0.006181 | 711.7752 | 0.091441 | 91.6309 | 647.4742 | 110.392 | 94.5905 | 61.584 | 21.602 | 75.7751 | 2069.0111 | -0.008702
648.1332 | 79.3053 | 67.9989 | 43.8913 | 17.275 | 51.8851 | 1493.7895 | 0.005833 | 717.4856 | 0.091462 | 91.8812 | 648.1332 | 111.0393 | 95.2087 | 61.4544 | 24.1876 | 72.6469 | 2091.5305 | 0.0081671
648.7923 | 80.2487 | 68.8157 | 44.0366 | 19.2577 | 50.0655 | 1512.3383 | 0.005816 | 723.7436 | 0.09131 | 92.1315 | 648.7923 | 111.8772 | 95.9381 | 61.3929 | 26.8477 | 69.7979 | 2108.3983 | 0.0081083
649.4514 | 81.1208 | 69.6677 | 44.1462 | 21.2155 | 48.3814 | 1522.9078 | 0.003292 | 729.7274 | 0.091273 | 92.3817 | 649.4514 | 112.5167 | 96.6309 | 61.232 | 29.4265 | 67.1063 | 2112.3142 | 0.0045661
650.1104 | 82.0928 | 70.4902 | 44.2756 | 23.1586 | 46.8653 | 1526.6297 | -0.000579 | 736.3727 | 0.091226 | 92.6319 | 650.1104 | 113.2003 | 97.2011 | 61.053 | 31.9341 | 64.624 | 2105.1171 | -0.0007984
650.7695 | 83.0041 | 71.3461 | 44.3629 | 24.9818 | 45.386 | 1522.4949 | 0.001467 | 742.2317 | 0.091256 | 92.8821 | 650.7695 | 113.8236 | 97.8368 | 60.8348 | 34.2575 | 62.2378 | 2087.7969 | 0.0020117
651.4285 | 83.9187 | 72.1387 | 44.4408 | 26.7286 | 44.0412 | 1511.2297 | -0.008717 | 748.3936 | 0.091154 | 93.1322 | 651.4285 | 114.5651 | 98.4831 | 60.6702 | 36.4897 | 60.1247 | 2063.1182 | -0.0119
652.0875 | 84.8923 | 72.9845 | 44.4907 | 28.3773 | 42.6926 | 1493.7081 | -0.003035 | 754.6271 | 0.091102 | 93.3821 | 652.0875 | 115.3107 | 99.1361 | 60.4325 | 38.5454 | 57.99 | 2028.9297 | -0.0041225
652.7466 | 85.7215 | 73.7641 | 44.4871 | 29.9328 | 41.4581 | 1468.9138 | -0.000303 | 759.5317 | 0.091122 | 93.6316 | 652.7466 | 115.9687 | 99.7921 | 60.1846 | 40.4948 | 56.0868 | 1987.2263 | -0.00040991
653.4056 | 86.5784 | 74.5719 | 44.4717 | 31.3933 | 40.2637 | 1439.3502 | -0.00462 | 765.4383 | 0.090914 | 93.8806 | 653.4056 | 116.8 | 100.6024 | 59.9953 | 42.3517 | 54.3184 | 1941.7784 | -0.0062327
654.0647 | 87.4401 | 75.3021 | 44.4181 | 32.7543 | 39.1569 | 1404.9588 | -0.000787 | 770.9111 | 0.090945 | 94.1291 | 654.0647 | 117.396 | 101.0997 | 59.6353 | 43.9755 | 52.5715 | 1886.2807 | -0.0010566
654.7238 | 88.2134 | 76.0644 | 44.3515 | 34.0226 | 38.0986 | 1366.8997 | -0.000809 | 775.6457 | 0.090906 | 94.3768 | 654.7238 | 118.0715 | 101.8104 | 59.3633 | 45.5384 | 50.994 | 1829.5617 | -0.0010828
655.3828 | 89.0519 | 76.8066 | 44.25 | 35.1873 | 37.1322 | 1326.0018 | -0.006479 | 780.2865 | 0.090949 | 94.6237 | 655.3828 | 118.7383 | 102.4109 | 59.0012 | 46.9173 | 49.5107 | 1768.0378 | -0.0086388
656.0419 | 89.8673 | 77.5522 | 44.1235 | 36.2916 | 36.2688 | 1283.4198 | -0.004196 | 784.9527 | 0.090743 | 94.8697 | 656.0419 | 119.6942 | 103.2917 | 58.768 | 48.3368 | 48.3064 | 1709.3858 | -0.0055886
656.7009 | 90.7213 | 78.3635 | 44.0011 | 37.3651 | 35.5602 | 1239.8252 | -0.008712 | 790.5713 | 0.09072 | 95.1146 | 656.7009 | 120.3136 | 103.9248 | 58.3537 | 49.5532 | 47.1596 | 1644.2421 | -0.011554
657.36 | 91.5914 | 79.1419 | 43.8571 | 38.406 | 34.9245 | 1195.953 | -0.001908 | 795.5768 | 0.090752 | 95.3582 | 657.36 | 120.9696 | 104.5269 | 57.9244 | 50.7248 | 46.1266 | 1579.5588 | -0.00252
658.019 | 92.4519 | 79.9232 | 43.6835 | 39.357 | 34.4385 | 1150.7853 | -0.002837 | 800.8832 | 0.090557 | 95.6006 | 658.019 | 121.867 | 105.3521 | 57.5821 | 51.8791 | 45.3957 | 1516.9265 | -0.0037396
658.678 | 93.3419 | 80.7153 | 43.4942 | 40.3073 | 34.074 | 1107.169 | 0.001664 | 807.1355 | 0.090589 | 95.8415 | 658.678 | 122.3514 | 105.8006 | 57.0116 | 52.8343 | 44.6637 | 1451.2627 | 0.0021811
659.3371 | 94.2462 | 81.5143 | 43.3024 | 41.2342 | 33.8342 | 1064.0171 | -0.007894 | 812.6973 | 0.090491 | 96.0809 | 659.3371 | 123.1305 | 106.4966 | 56.5736 | 53.8716 | 44.2036 | 1390.1144 | -0.010313
659.9962 | 95.1874 | 82.3791 | 43.1 | 42.1935 | 33.6377 | 1022.0296 | 8e-06 | 818.6869 | 0.090421 | 96.3186 | 659.9962 | 123.8522 | 107.1867 | 56.0791 | 54.8996 | 43.7673 | 1329.8033 | 1.0409e-05
660.6552 | 96.1375 | 83.2332 | 42.8759 | 43.1339 | 33.5595 | 981.0074 | -0.002514 | 824.6065 | 0.090386 | 96.5546 | 660.6552 | 124.5422 | 107.8252 | 55.5439 | 55.8782 | 43.4749 | 1270.8546 | -0.0032568
661.3143 | 97.0749 | 84.0961 | 42.6466 | 44.1091 | 33.5141 | 941.8767 | 0.002677 | 831.0828 | 0.090324 | 96.789 | 661.3143 | 125.1658 | 108.4314 | 54.9875 | 56.8731 | 43.2122 | 1214.431 | 0.0034517
661.9733 | 98.0521 | 84.9998 | 42.3933 | 45.051 | 33.5907 | 904.0209 | -0.0046 | 837.3356 | 0.090413 | 97.0221 | 661.9733 | 125.6606 | 108.9332 | 54.3299 | 57.736 | 43.0489 | 1158.566 | -0.0058952
662.6324 | 99.0186 | 85.8374 | 42.1296 | 46.0323 | 33.6746 | 867.1327 | -0.002117 | 843.6348 | 0.090215 | 97.254 | 662.6324 | 126.5297 | 109.6863 | 53.8347 | 58.8218 | 43.0307 | 1108.0545 | -0.0027052
663.2914 | 100.0409 | 86.734 | 41.8787 | 47.0567 | 33.8849 | 832.1326 | 0.003323 | 850.6788 | 0.090291 | 97.4851 | 663.2914 | 126.9707 | 110.0817 | 53.1519 | 59.7238 | 43.0063 | 1056.1325 | 0.0042175
663.9505 | 101.048 | 87.6632 | 41.5996 | 48.1601 | 34.1949 | 798.838 | 0.003549 | 857.7343 | 0.090205 | 97.7155 | 663.9505 | 127.6165 | 110.7124 | 52.5373 | 60.8228 | 43.1857 | 1008.876 | 0.0044821
664.6095 | 102.0961 | 88.6105 | 41.3433 | 49.2604 | 34.5786 | 766.4289 | -0.000524 | 864.6013 | 0.090175 | 97.9455 | 664.6095 | 128.2599 | 111.3184 | 51.9382 | 61.8842 | 43.4399 | 962.8385 | -0.00065828
665.2686 | 103.1376 | 89.5771 | 41.0514 | 50.3928 | 35.0878 | 734.9278 | -0.002323 | 872.0132 | 0.090154 | 98.1753 | 665.2686 | 128.7993 | 111.8647 | 51.2654 | 62.931 | 43.8179 | 917.7847 | -0.002901
665.9276 | 104.1987 | 90.5154 | 40.7626 | 51.6109 | 35.7049 | 704.8972 | -3.6e-05 | 879.2419 | 0.090028 | 98.4051 | 665.9276 | 129.5368 | 112.5261 | 50.6748 | 64.1612 | 44.3873 | 876.3075 | -4.4754e-05
666.5867 | 105.2933 | 91.4675 | 40.4711 | 52.8367 | 36.4613 | 675.8584 | -0.000703 | 886.7454 | 0.090076 | 98.6352 | 666.5867 | 130.0246 | 112.9513 | 49.9769 | 65.247 | 45.0253 | 834.6038 | -0.00086812
667.2457 | 106.2774 | 92.3696 | 40.1469 | 54.0851 | 37.3134 | 647.3925 | -0.002732 | 893.1522 | 0.090042 | 98.8659 | 667.2457 | 130.6526 | 113.555 | 49.3548 | 66.4897 | 45.8714 | 795.8749 | -0.0033586
667.9048 | 107.3274 | 93.3111 | 39.826 | 55.4112 | 38.3063 | 619.9708 | -0.000227 | 900.8008 | 0.089898 | 99.0972 | 667.9048 | 131.3389 | 114.1868 | 48.7359 | 67.8078 | 46.8763 | 758.6716 | -0.00027778
668.5638 | 108.384 | 94.226 | 39.5078 | 56.7671 | 39.4086 | 593.5083 | -0.00205 | 907.6332 | 0.08971 | 99.3294 | 668.5638 | 132.2186 | 114.9471 | 48.196 | 69.2507 | 48.0749 | 724.0261 | -0.0025008
669.2229 | 109.4303 | 95.1585 | 39.1646 | 58.1686 | 40.6572 | 567.889 | 0.000431 | 914.5315 | 0.090045 | 99.5629 | 669.2229 | 132.305 | 115.0498 | 47.3513 | 70.3279 | 49.1559 | 686.5973 | 0.00052109
669.882 | 110.3352 | 96.05 | 38.8233 | 59.5913 | 41.9813 | 542.8575 | 0.006082 | 921.1466 | 0.089758 | 99.7978 | 669.882 | 133.178 | 115.9353 | 46.8609 | 71.9285 | 50.6728 | 655.2459 | 0.0073412
670.541 | 111.338 | 96.9378 | 38.4774 | 61.0731 | 43.4189 | 518.6291 | 0.006745 | 928.1094 | 0.089709 | 100.0342 | 670.541 | 133.7695 | 116.468 | 46.2296 | 73.3776 | 52.1666 | 623.1182 | 0.0081039
671.2 | 112.388 | 97.8258 | 38.149 | 62.6152 | 44.9521 | 495.4857 | -0.00097 | 934.7438 | 0.089777 | 100.2722 | 671.2 | 134.2894 | 116.8894 | 45.5833 | 74.8172 | 53.7121 | 592.0425 | -0.001159
671.8591 | 113.3836 | 98.7091 | 37.8054 | 64.1829 | 46.5457 | 472.563 | 0.004931 | 941.1985 | 0.089689 | 100.5115 | 671.8591 | 135.0038 | 117.5311 | 45.0142 | 76.4214 | 55.4212 | 562.6721 | 0.0058713
672.5181 | 114.3255 | 99.5659 | 37.462 | 65.7873 | 48.2716 | 450.7751 | -0.001632 | 947.7083 | 0.089558 | 100.7519 | 672.5181 | 135.7117 | 118.1912 | 44.4698 | 78.0938 | 57.3015 | 535.0994 | -0.0019373
673.1772 | 115.1533 | 100.3064 | 37.0775 | 67.3716 | 50.0626 | 428.9999 | 0.003823 | 952.8621 | 0.089664 | 100.9933 | 673.1772 | 136.1189 | 118.5689 | 43.8281 | 79.6377 | 59.1773 | 507.1066 | 0.004519
673.8362 | 115.975 | 101.0179 | 36.6942 | 68.9546 | 51.9172 | 408.1992 | 0.002458 | 957.7786 | 0.08969 | 101.2354 | 673.8362 | 136.6749 | 119.0482 | 43.2436 | 81.262 | 61.1836 | 481.0568 | 0.0028967
674.4953 | 116.7688 | 101.6656 | 36.2761 | 70.545 | 53.8614 | 387.7669 | -0.006182 | 961.5891 | 0.089562 | 101.4781 | 674.4953 | 137.5896 | 119.7933 | 42.7444 | 83.1237 | 63.4653 | 456.9088 | -0.0072843
675.1544 | 117.4161 | 102.2611 | 35.8478 | 72.1383 | 55.8715 | 367.9246 | -0.000325 | 965.2033 | 0.089463 | 101.7211 | 675.1544 | 138.317 | 120.4643 | 42.229 | 84.9795 | 65.817 | 433.4179 | -0.00038285
675.8134 | 118.0646 | 102.8263 | 35.3865 | 73.7179 | 57.9023 | 348.8661 | -0.008987 | 968.7328 | 0.089466 | 101.9643 | 675.8134 | 138.9014 | 120.9738 | 41.6318 | 86.7281 | 68.1213 | 410.4361 | -0.010573
676.4725 | 118.6554 | 103.3302 | 34.9027 | 75.2994 | 60.0746 | 330.4388 | 0.002534 | 971.1668 | 0.089325 | 102.2074 | 676.4725 | 139.7983 | 121.7424 | 41.122 | 88.7168 | 70.7791 | 389.3189 | 0.0029855
677.1315 | 119.3148 | 103.9228 | 34.4307 | 76.9453 | 62.2792 | 313.3317 | 0.001818 | 974.5468 | 0.089353 | 102.4503 | 677.1315 | 140.3776 | 122.2685 | 40.5088 | 90.5286 | 73.2735 | 368.6445 | 0.0021389
677.7905 | 120.1039 | 104.5046 | 33.9399 | 78.6001 | 64.6681 | 296.8944 | 0.009285 | 978.2439 | 0.089343 | 102.6928 | 677.7905 | 141.12 | 122.7911 | 39.8788 | 92.3537 | 75.9838 | 348.8456 | 0.01091
678.4496 | 120.8665 | 105.1845 | 33.4408 | 80.342 | 67.0681 | 281.8542 | 0.008605 | 982.0025 | 0.089161 | 102.9346 | 678.4496 | 142.0949 | 123.6587 | 39.3142 | 94.453 | 78.8477 | 331.3579 | 0.010116
679.1086 | 121.6874 | 105.9126 | 32.9284 | 82.1155 | 69.592 | 267.2075 | 0.008614 | 986.5537 | 0.089244 | 103.1756 | 679.1086 | 142.6009 | 124.115 | 38.5875 | 96.2281 | 81.5523 | 313.1306 | 0.010094
679.7677 | 122.6628 | 106.7608 | 32.4318 | 83.9715 | 72.2853 | 253.8586 | 0.010881 | 993.2566 | 0.08924 | 103.4157 | 679.7677 | 143.1122 | 124.5591 | 37.8386 | 97.9705 | 84.3362 | 296.1799 | 0.012695
680.4268 | 123.6746 | 107.6276 | 31.8921 | 85.8764 | 75.0233 | 241.1076 | 0.014035 | 999.5194 | 0.089119 | 103.6545 | 680.4268 | 143.9161 | 125.2428 | 37.1118 | 99.9316 | 87.3022 | 280.5691 | 0.016332
681.0858 | 124.6721 | 108.4584 | 31.3444 | 87.7294 | 77.8488 | 229.2516 | 0.00629 | 1005.7571 | 0.089025 | 103.8921 | 681.0858 | 144.6601 | 125.8469 | 36.3697 | 101.7947 | 90.3299 | 266.0063 | 0.0072984
681.7449 | 125.7388 | 109.3378 | 30.7726 | 89.6316 | 80.7036 | 217.828 | 0.004658 | 1012.592 | 0.088988 | 104.1285 | 681.7449 | 145.3022 | 126.3494 | 35.5604 | 103.5772 | 93.2601 | 251.7193 | 0.0053827
682.4039 | 126.91 | 110.2874 | 30.2076 | 91.5044 | 83.6527 | 207.0445 | 0.004842 | 1019.3736 | 0.089036 | 104.3636 | 682.4039 | 145.9298 | 126.816 | 34.7347 | 105.2181 | 96.1896 | 238.074 | 0.0055677
683.063 | 127.977 | 111.1941 | 29.6153 | 93.3726 | 86.6113 | 196.9431 | 0.00617 | 1027.3126 | 0.088987 | 104.5977 | 683.063 | 146.4278 | 127.2252 | 33.885 | 106.8344 | 99.0982 | 225.3369 | 0.0070595
683.722 | 129.0485 | 112.1063 | 29.0219 | 95.1798 | 89.5949 | 186.9734 | 0.009098 | 1034.4419 | 0.088937 | 104.8306 | 683.722 | 147.0456 | 127.7407 | 33.0693 | 108.4535 | 102.0898 | 213.0487 | 0.010367
684.381 | 130.1893 | 113.0413 | 28.4159 | 97.03 | 92.6278 | 177.6519 | 0.00319 | 1041.7096 | 0.088908 | 105.0624 | 684.381 | 147.6849 | 128.2325 | 32.2346 | 110.0695 | 105.0758 | 201.5259 | 0.0036187
685.0401 | 131.272 | 113.9345 | 27.8275 | 98.7964 | 95.5572 | 168.8748 | 0.003627 | 1049.3958 | 0.088873 | 105.2932 | 685.0401 | 148.205 | 128.6311 | 31.4171 | 111.5403 | 107.8833 | 190.6582 | 0.0040949
685.6992 | 132.3544 | 114.8357 | 27.2387 | 100.5309 | 98.6269 | 160.4182 | 0.001338 | 1056.4975 | 0.088835 | 105.523 | 685.6992 | 148.8095 | 129.1128 | 30.6252 | 113.0296 | 110.8888 | 180.3624 | 0.0015043
686.3582 | 133.4193 | 115.6732 | 26.657 | 102.148 | 101.6037 | 152.4194 | 0.006072 | 1063.3035 | 0.088746 | 105.7518 | 686.3582 | 149.5208 | 129.6331 | 29.8741 | 114.4757 | 113.8656 | 170.814 | 0.0068048
687.0173 | 134.486 | 116.6005 | 26.0949 | 103.8667 | 104.6745 | 145.1328 | -0.000764 | 1070.9142 | 0.088657 | 105.9796 | 687.0173 | 150.1175 | 130.1531 | 29.128 | 115.9392 | 116.841 | 162.0018 | -0.0008528
687.6763 | 135.6415 | 117.4935 | 25.5362 | 105.4485 | 107.6268 | 138.1974 | 0.003297 | 1077.6989 | 0.088722 | 106.2066 | 687.6763 | 150.6655 | 130.5073 | 28.3646 | 117.1283 | 119.5479 | 153.5045 | 0.0036622
688.3354 | 136.7227 | 118.3223 | 24.9885 | 107.1027 | 110.629 | 131.4461 | 0.003297 | 1084.8192 | 0.088745 | 106.4327 | 688.3354 | 151.1518 | 130.8095 | 27.6257 | 118.4058 | 122.3043 | 145.3183 | 0.003645
688.9944 | 137.8115 | 119.1494 | 24.438 | 108.5597 | 113.592 | 125.0625 | -0.002583 | 1091.6722 | 0.088664 | 106.6579 | 688.9944 | 151.8585 | 131.2942 | 26.9289 | 119.625 | 125.1703 | 137.81 | -0.0028463
689.6535 | 138.8188 | 119.992 | 23.8846 | 110.0649 | 116.589 | 119.0362 | -0.003028 | 1098.8392 | 0.088751 | 106.8823 | 689.6535 | 152.1403 | 131.5068 | 26.1766 | 120.6271 | 127.7773 | 130.4593 | -0.0033186
690.3126 | 139.9566 | 120.8674 | 23.3549 | 111.6166 | 119.5247 | 113.2385 | 0.003534 | 1106.3219 | 0.088536 | 107.106 | 690.3126 | 153.0396 | 132.166 | 25.5381 | 122.0504 | 130.6977 | 123.8239 | 0.0038644
690.9716 | 141.0047 | 121.7003 | 22.8101 | 113.0212 | 122.4077 | 107.8264 | 0.003264 | 1113.2841 | 0.088573 | 107.3288 | 690.9716 | 153.4768 | 132.4648 | 24.8276 | 123.018 | 133.2348 | 117.3637 | 0.0035527
691.6306 | 142.075 | 122.4797 | 22.2541 | 114.3782 | 125.2484 | 102.5467 | 0.008871 | 1119.476 | 0.088581 | 107.551 | 691.6306 | 154.0904 | 132.838 | 24.1362 | 124.0512 | 135.8407 | 111.2191 | 0.0096212
692.2897 | 143.003 | 123.2294 | 21.6946 | 115.7755 | 127.9945 | 97.5311 | 0.000654 | 1126.3668 | 0.088463 | 107.7723 | 692.2897 | 154.6712 | 133.2842 | 23.4648 | 125.2221 | 138.4381 | 105.4891 | 0.00070736
692.9487 | 144.1824 | 124.0693 | 21.1425 | 117.1178 | 130.8101 | 92.8842 | -0.002324 | 1133.2955 | 0.088416 | 107.9928 | 692.9487 | 155.3943 | 133.7171 | 22.7866 | 126.225 | 140.9821 | 100.107 | -0.0025047
693.6078 | 145.2045 | 124.8182 | 20.5776 | 118.4188 | 133.5451 | 88.4872 | -0.002789 | 1140.1428 | 0.088447 | 108.2126 | 693.6078 | 155.8175 | 133.9411 | 22.0816 | 127.074 | 143.3058 | 94.9547 | -0.0029928
694.2668 | 146.2561 | 125.6906 | 20.0207 | 119.7219 | 136.367 | 84.3404 | -7e-05 | 1147.0768 | 0.088389 | 108.4315 | 694.2668 | 156.4156 | 134.4216 | 21.4115 | 128.0383 | 145.8396 | 90.199 | -7.4863e-05
694.9259 | 147.3084 | 126.4392 | 19.4536 | 121.0265 | 139.0126 | 80.4142 | 0.006952 | 1153.6327 | 0.088427 | 108.6496 | 694.9259 | 156.8933 | 134.6662 | 20.7194 | 128.9013 | 148.0577 | 85.6464 | 0.0074043
695.585 | 148.4249 | 127.1875 | 18.8875 | 122.1958 | 141.6217 | 76.4543 | 0.00353 | 1160.8284 | 0.088365 | 108.8669 | 695.585 | 157.5271 | 134.9874 | 20.0457 | 129.6895 | 150.3067 | 81.1429 | 0.0037465
696.244 | 149.4718 | 127.9481 | 18.319 | 123.4043 | 144.2302 | 72.7409 | 0.004404 | 1167.6548 | 0.088419 | 109.0833 | 696.244 | 157.9267 | 135.1855 | 19.3552 | 130.3847 | 152.3886 | 76.8555 | 0.0046531
696.9006 | 150.4326 | 128.6441 | 17.7451 | 124.5569 | 146.6652 | 69.2301 | 0.004819 | 1173.7352 | 0.088368 | 109.298 | 696.9006 | 158.5216 | 135.5615 | 18.6993 | 131.2545 | 154.5516 | 72.9527 | 0.0050781
697.5572 | 151.5418 | 129.4742 | 17.1995 | 125.8105 | 149.2185 | 65.9334 | 0.009993 | 1181.9613 | 0.088227 | 109.5119 | 697.5572 | 159.1429 | 135.9685 | 18.0622 | 132.121 | 156.7031 | 69.2405 | 0.010494
698.2138 | 152.6201 | 130.1855 | 16.6423 | 126.9507 | 151.5914 | 62.8196 | 0.000656 | 1188.5766 | 0.088172 | 109.725 | 698.2138 | 159.7934 | 136.3044 | 17.4245 | 132.9176 | 158.7164 | 65.7721 | 0.00068683
698.8704 | 153.6413 | 130.8656 | 16.099 | 128.0353 | 153.912 | 59.8646 | 0.00629 | 1194.8425 | 0.088188 | 109.9371 | 698.8704 | 160.3 | 136.5373 | 16.7967 | 133.5842 | 160.5824 | 62.4591 | 0.0065626
699.527 | 154.7278 | 131.5072 | 15.563 | 129.1652 | 156.1565 | 56.892 | 0.003763 | 1201.3289 | 0.088063 | 110.1484 | 699.527 | 161.0985 | 136.9218 | 16.2038 | 134.4834 | 162.586 | 59.2345 | 0.0039179
700.1835 | 155.7086 | 132.2401 | 15.0358 | 130.2452 | 158.3697 | 54.3405 | 0.008488 | 1207.6482 | 0.088224 | 110.3587 | 700.1835 | 161.2839 | 136.9751 | 15.5742 | 134.9088 | 164.0403 | 56.2862 | 0.0087919
700.8401 | 156.8286 | 132.9358 | 14.5226 | 131.3396 | 160.5401 | 51.7422 | 0.001208 | 1214.2539 | 0.088155 | 110.5682 | 700.8401 | 161.9932 | 137.3136 | 15.0008 | 135.6649 | 165.827 | 53.4462 | 0.0012478
701.4968 | 157.7933 | 133.5661 | 14.0252 | 132.34 | 162.5071 | 49.3448 | 0.004135 | 1220.651 | 0.088112 | 110.7767 | 701.4968 | 162.5215 | 137.5683 | 14.4454 | 136.3055 | 167.3765 | 50.8234 | 0.0042589
702.1533 | 158.8383 | 134.2264 | 13.53 | 133.3505 | 164.5555 | 47.1168 | 0.002523 | 1227.1927 | 0.088108 | 110.9843 | 702.1533 | 163.0377 | 137.7751 | 13.8877 | 136.876 | 168.906 | 48.3625 | 0.0025897
702.8099 | 159.8972 | 134.8599 | 13.0622 | 134.3844 | 166.4982 | 45.0427 | 0.003856 | 1233.9284 | 0.088051 | 111.1911 | 702.8099 | 163.639 | 138.0159 | 13.3679 | 137.5292 | 170.3945 | 46.0967 | 0.0039462
703.4666 | 160.9309 | 135.4919 | 12.6088 | 135.4006 | 168.3863 | 42.9325 | 0.003824 | 1240.6882 | 0.088037 | 111.397 | 703.4666 | 164.1292 | 138.1847 | 12.8594 | 138.0914 | 171.7327 | 43.7857 | 0.0039
704.1231 | 161.9358 | 136.0614 | 12.1796 | 136.382 | 170.2061 | 41.0504 | 0.002417 | 1247.576 | 0.088076 | 111.602 | 704.1231 | 164.4719 | 138.1922 | 12.3703 | 138.5179 | 172.8717 | 41.6933 | 0.0024549
704.7797 | 162.9951 | 136.6665 | 11.767 | 137.3361 | 172.0345 | 39.2064 | -0.00103 | 1253.7051 | 0.087971 | 111.8062 | 704.7797 | 165.237 | 138.5462 | 11.9288 | 139.225 | 174.4006 | 39.7456 | -0.0010442
705.4363 | 164.0465 | 137.2587 | 11.3751 | 138.3311 | 173.68 | 37.5455 | -0.003112 | 1260.1484 | 0.087911 | 112.0095 | 705.4363 | 165.866 | 138.781 | 11.5013 | 139.8653 | 175.6063 | 37.9619 | -0.0031465
706.0929 | 165.1877 | 137.9569 | 11.0101 | 139.3542 | 175.4825 | 35.9227 | -0.000663 | 1267.5819 | 0.08807 | 112.212 | 706.0929 | 166.0408 | 138.6694 | 11.067 | 140.0739 | 176.3888 | 36.1082 | -0.00066642
706.7495 | 166.1078 | 138.4277 | 10.6574 | 140.2529 | 177.0084 | 34.3846 | 0.003372 | 1274.2593 | 0.087862 | 112.4137 | 706.7495 | 166.7831 | 138.9904 | 10.7007 | 140.823 | 177.728 | 34.5243 | 0.0033857
707.4061 | 167.1449 | 138.9979 | 10.3277 | 141.2172 | 178.5343 | 32.9313 | -0.004833 | 1280.8538 | 0.087878 | 112.6145 | 707.4061 | 167.2284 | 139.0673 | 10.3328 | 141.2878 | 178.6235 | 32.9477 | -0.0048354
708.0627 | 168.2119 | 139.5552 | 10.016 | 142.1007 | 180.0954 | 31.5525 | 0.008499 | 1287.5846 | 0.08786 | 112.8145 | 708.0627 | 167.7463 | 139.169 | 9.9882 | 141.7074 | 179.5969 | 31.4651 | 0.0084755
708.7193 | 169.2298 | 140.0582 | 9.7244 | 143.0731 | 181.4275 | 30.2815 | 0.001432 | 1294.5251 | 0.087846 | 113.0137 | 708.7193 | 168.1805 | 139.1898 | 9.6641 | 142.186 | 180.3026 | 30.0937 | 0.0014231
709.3759 | 170.3345 | 140.5656 | 9.4573 | 143.9474 | 182.7799 | 29.0233 | 0.010236 | 1300.9493 | 0.087789 | 113.2121 | 709.3759 | 168.8473 | 139.3383 | 9.3747 | 142.6906 | 181.184 | 28.7699 | 0.010147
710.0325 | 171.3151 | 141.0314 | 9.2025 | 144.8144 | 184.0905 | 27.8602 | 0.007915 | 1306.937 | 0.087764 | 113.4098 | 710.0325 | 169.3855 | 139.4429 | 9.0988 | 143.1834 | 182.017 | 27.5464 | 0.0078259
710.6891 | 172.3188 | 141.4463 | 8.9708 | 145.6328 | 185.2744 | 26.653 | 0.012157 | 1313.4359 | 0.08773 | 113.6066 | 710.6891 | 169.8945 | 139.4564 | 8.8446 | 143.5839 | 182.6679 | 26.278 | 0.011986
711.3456 | 173.3544 | 141.933 | 8.7585 | 146.4994 | 186.4874 | 25.6808 | 0.007995 | 1320.495 | 0.087746 | 113.8029 | 711.3456 | 170.264 | 139.4028 | 8.6024 | 143.8878 | 183.163 | 25.223 | 0.0078525
712.0023 | 174.4043 | 142.3533 | 8.5596 | 147.3699 | 187.6371 | 24.656 | -0.001163 | 1326.8172 | 0.087698 | 113.9986 | 712.0023 | 170.8659 | 139.4652 | 8.3859 | 144.38 | 183.8303 | 24.1558 | -0.0011394
712.6588 | 175.345 | 142.7077 | 8.3767 | 148.1671 | 188.7 | 23.6941 | 0.007132 | 1333.4208 | 0.087617 | 114.1941 | 712.6588 | 171.3887 | 139.4878 | 8.1877 | 144.824 | 184.4424 | 23.1595 | 0.0069711
713.3154 | 176.3542 | 143.1249 | 8.2101 | 148.9507 | 189.6421 | 22.7521 | -0.000666 | 1339.99 | 0.087549 | 114.3893 | 713.3154 | 171.9572 | 139.5564 | 8.0054 | 145.237 | 184.9138 | 22.1848 | -0.00064939
713.972 | 177.3687 | 143.4321 | 8.061 | 149.7683 | 190.5969 | 21.9509 | -0.000284 | 1345.8729 | 0.087685 | 114.5845 | 713.972 | 172.2159 | 139.2653 | 7.8269 | 145.4173 | 185.0598 | 21.3132 | -0.00027575
714.6286 | 178.4588 | 143.8379 | 7.9254 | 150.571 | 191.5852 | 21.106 | 0.001485 | 1352.8433 | 0.087574 | 114.7798 | 714.6286 | 172.8943 | 139.3529 | 7.6783 | 145.8762 | 185.6115 | 20.4479 | 0.0014387
715.2852 | 179.4359 | 144.1061 | 7.7998 | 151.3453 | 192.4743 | 20.337 | -0.003072 | 1359.1461 | 0.087535 | 114.9754 | 715.2852 | 173.4075 | 139.2646 | 7.5377 | 146.2606 | 186.0078 | 19.6538 | -0.0029688
715.9418 | 180.3636 | 144.3411 | 7.6865 | 152.05 | 193.1868 | 19.6078 | 0.002128 | 1364.0158 | 0.087562 | 115.1713 | 715.9418 | 173.9243 | 139.1879 | 7.412 | 146.6216 | 186.2897 | 18.9078 | 0.002052
716.5984 | 181.3853 | 144.676 | 7.5818 | 152.8223 | 194.0652 | 18.9319 | 0.003019 | 1371.5751 | 0.087466 | 115.3678 | 716.5984 | 174.4325 | 139.1303 | 7.2911 | 146.9644 | 186.6263 | 18.2062 | 0.0029033
717.255 | 182.4155 | 144.9089 | 7.4935 | 153.5571 | 194.8065 | 18.2476 | -0.002907 | 1377.591 | 0.087353 | 115.565 | 717.255 | 175.1813 | 139.1621 | 7.1963 | 147.4674 | 187.0809 | 17.5239 | -0.0027917
717.9116 | 183.4616 | 145.1876 | 7.4173 | 154.3278 | 195.4601 | 17.6495 | 0.002477 | 1384.7285 | 0.087394 | 115.763 | 717.9116 | 175.4961 | 138.8839 | 7.0953 | 147.6273 | 186.9738 | 16.8832 | 0.0023695
718.5682 | 184.4015 | 145.3733 | 7.3563 | 154.9712 | 196.0306 | 17.0368 | -0.0035 | 1390.634 | 0.087456 | 115.962 | 718.5682 | 175.8238 | 138.6111 | 7.0141 | 147.7625 | 186.9119 | 16.2443 | -0.0033372
719.2248 | 185.4169 | 145.5133 | 7.304 | 155.6792 | 196.752 | 16.4701 | 0.004345 | 1396.689 | 0.087348 | 116.1621 | 719.2248 | 176.5477 | 138.5528 | 6.9546 | 148.2324 | 187.3405 | 15.6823 | 0.0041372
719.8813 | 186.3532 | 145.7176 | 7.2616 | 156.3074 | 197.2668 | 15.9215 | 0.003123 | 1402.9382 | 0.087443 | 116.3635 | 719.8813 | 176.7619 | 138.2177 | 6.8878 | 148.2625 | 187.1138 | 15.102 | 0.0029623
720.538 | 187.3119 | 145.8456 | 7.2314 | 157.048 | 197.8389 | 15.405 | -0.001264 | 1408.873 | 0.08728 | 116.5662 | 720.538 | 177.5623 | 138.2543 | 6.855 | 148.8737 | 187.5415 | 14.6032 | -0.0011982
721.1946 | 188.3553 | 145.8953 | 7.2161 | 157.6478 | 198.2931 | 14.9482 | -0.000421 | 1415.5269 | 0.087283 | 116.7703 | 721.1946 | 178.0181 | 137.8883 | 6.82 | 148.9958 | 187.4104 | 14.1278 | -0.00039789
721.8511 | 189.2986 | 145.9574 | 7.2082 | 158.2521 | 198.6642 | 14.4796 | 0.004737 | 1421.5145 | 0.087404 | 116.9752 | 721.8511 | 178.2205 | 137.4157 | 6.7864 | 148.9909 | 187.038 | 13.6323 | 0.0044598
722.5078 | 190.2626 | 146.0335 | 7.2113 | 158.8757 | 199.0596 | 13.9806 | -7.9e-05 | 1427.6881 | 0.087232 | 117.181 | 722.5078 | 179.0193 | 137.4039 | 6.7852 | 149.4872 | 187.2965 | 13.1545 | -7.4332e-05
723.1644 | 191.2776 | 146.1275 | 7.2275 | 159.5332 | 199.4545 | 13.573 | 0.001002 | 1434.4842 | 0.087177 | 117.3872 | 723.1644 | 179.5516 | 137.1693 | 6.7844 | 149.7532 | 187.2272 | 12.741 | 0.00094057
723.8209 | 192.247 | 146.1443 | 7.2565 | 160.1503 | 199.8593 | 13.1781 | -0.001085 | 1439.895 | 0.08713 | 117.5936 | 723.8209 | 180.1963 | 136.9834 | 6.8017 | 150.1115 | 187.3314 | 12.3521 | -0.001017
724.4775 | 193.3242 | 146.1603 | 7.2923 | 160.7687 | 200.1693 | 12.8265 | 0.004545 | 1446.4729 | 0.087185 | 117.8 | 724.4775 | 180.5836 | 136.5279 | 6.8117 | 150.1736 | 186.9776 | 11.9812 | 0.0042455
725.1342 | 194.2517 | 146.0908 | 7.3325 | 161.271 | 200.3993 | 12.4485 | 0.002249 | 1452.8775 | 0.087162 | 118.0061 | 725.1342 | 181.0143 | 136.1354 | 6.8328 | 150.2811 | 186.743 | 11.6002 | 0.0020957
725.7907 | 195.2389 | 146.0577 | 7.3834 | 161.8501 | 200.7306 | 12.1092 | -0.001633 | 1458.9505 | 0.087135 | 118.2117 | 725.7907 | 181.5492 | 135.8165 | 6.8657 | 150.5015 | 186.6558 | 11.2601 | -0.0015185
726.4473 | 196.1476 | 145.8888 | 7.4358 | 162.3491 | 200.9435 | 11.7576 | -0.003716 | 1464.6644 | 0.087049 | 118.4165 | 726.4473 | 182.1769 | 135.4978 | 6.9062 | 150.7856 | 186.6312 | 10.9202 | -0.0034513
727.1039 | 197.0976 | 145.8079 | 7.494 | 162.9115 | 201.0426 | 11.4517 | -0.002199 | 1471.2746 | 0.087042 | 118.6202 | 727.1039 | 182.5659 | 135.0577 | 6.9415 | 150.9003 | 186.22 | 10.6074 | -0.0020369
727.7605 | 197.9914 | 145.5907 | 7.5589 | 163.3461 | 201.2755 | 11.1518 | 0.002473 | 1476.1297 | 0.08703 | 118.8227 | 727.7605 | 183.126 | 134.6596 | 6.9914 | 151.0819 | 186.1635 | 10.3145 | 0.0022873
728.4171 | 199.0046 | 145.4118 | 7.6148 | 163.8318 | 201.4535 | 10.847 | 0.004855 | 1482.3064 | 0.086976 | 119.0236 | 728.4171 | 183.722 | 134.2449 | 7.0301 | 151.2503 | 185.9828 | 10.014 | 0.0044822
729.0737 | 199.9413 | 145.245 | 7.6774 | 164.3832 | 201.6024 | 10.5931 | -0.001173 | 1488.9096 | 0.086924 | 119.2227 | 729.0737 | 184.1839 | 133.7983 | 7.0723 | 151.4281 | 185.7141 | 9.7582 | -0.0010806
729.7303 | 201.2018 | 145.2529 | 7.7411 | 165.0889 | 201.9988 | 10.3048 | 0.003225 | 1496.1691 | 0.08701 | 119.4197 | 729.7303 | 184.5692 | 133.2454 | 7.1012 | 151.4416 | 185.3003 | 9.453 | 0.0029584
730.3869 | 202.458 | 145.1615 | 7.797 | 165.7699 | 202.3733 | 10.0843 | -0.004159 | 1504.522 | 0.086775 | 119.6145 | 730.3869 | 185.4922 | 132.9971 | 7.1436 | 151.8786 | 185.4147 | 9.2393 | -0.0038105
731.0435 | 203.949 | 145.2847 | 7.8556 | 166.6557 | 203.0353 | 9.8691 | -0.003137 | 1514.3285 | 0.086911 | 119.8069 | 731.0435 | 185.6556 | 132.2533 | 7.151 | 151.7074 | 184.8239 | 8.9838 | -0.0028556
731.7001 | 205.3885 | 145.3306 | 7.9021 | 167.4777 | 203.5814 | 9.6775 | -0.001637 | 1524.1396 | 0.086912 | 119.9973 | 731.7001 | 186.0567 | 131.6517 | 7.1584 | 151.7142 | 184.4197 | 8.7666 | -0.0014829
732.3566 | 207.0047 | 145.4374 | 7.9422 | 168.3795 | 204.1961 | 9.4823 | 0.000782 | 1535.2364 | 0.086904 | 120.1857 | 732.3566 | 186.4728 | 131.0121 | 7.1544 | 151.6787 | 183.9428 | 8.5418 | 0.00070444
733.0132 | 208.5947 | 145.5287 | 7.9706 | 169.3427 | 204.8904 | 9.3006 | 0.002971 | 1546.0487 | 0.086731 | 120.3724 | 733.0132 | 187.2542 | 130.6403 | 7.1552 | 152.018 | 183.9289 | 8.3491 | 0.002667
733.6699 | 210.3091 | 145.6161 | 7.9836 | 170.2768 | 205.5979 | 9.1068 | 0.002218 | 1557.4517 | 0.086751 | 120.5575 | 733.6699 | 187.6561 | 129.9313 | 7.1237 | 151.9358 | 183.4523 | 8.1258 | 0.0019791
734.3264 | 212.2179 | 145.8485 | 7.9839 | 171.3829 | 206.4073 | 8.9503 | -0.008513 | 1570.0519 | 0.086682 | 120.7412 | 734.3264 | 188.2756 | 129.394 | 7.0832 | 152.0477 | 183.1206 | 7.9406 | -0.0075526
734.983 | 214.2898 | 146.0862 | 7.9802 | 172.6017 | 207.391 | 8.7937 | -0.00519 | 1583.8393 | 0.086704 | 120.9238 | 734.983 | 188.6959 | 128.6382 | 7.027 | 151.9868 | 182.621 | 7.7434 | -0.0045701
735.6396 | 216.3221 | 146.3972 | 7.9561 | 173.8498 | 208.3469 | 8.6586 | 0.0013 | 1597.8346 | 0.086676 | 121.1054 | 735.6396 | 189.1623 | 128.0167 | 6.9572 | 152.0225 | 182.1884 | 7.5714 | 0.0011368
736.2962 | 218.6523 | 146.7239 | 7.9158 | 175.2661 | 209.4298 | 8.5454 | -0.001524 | 1614.5444 | 0.086655 | 121.2862 | 736.2962 | 189.5494 | 127.1948 | 6.8622 | 151.938 | 181.5544 | 7.408 | -0.0013212
736.9528 | 221.0249 | 147.1345 | 7.8694 | 176.6334 | 210.6256 | 8.448 | 0.000393 | 1630.6782 | 0.086661 | 121.4664 | 736.9528 | 189.9778 | 126.4667 | 6.764 | 151.8219 | 181.0393 | 7.2613 | 0.0003378
737.6094 | 223.6124 | 147.5019 | 7.8078 | 178.161 | 211.9388 | 8.3227 | 0.005467 | 1648.482 | 0.086724 | 121.6462 | 737.6094 | 190.2697 | 125.508 | 6.6436 | 151.5956 | 180.3368 | 7.0817 | 0.0046518
738.266 | 226.1643 | 147.9525 | 7.7426 | 179.7003 | 213.2045 | 8.1929 | 0.003725 | 1667.0471 | 0.086691 | 121.8257 | 738.266 | 190.6511 | 124.7204 | 6.5268 | 151.483 | 179.7263 | 6.9064 | 0.0031401
738.918 | 228.7833 | 148.3355 | 7.6531 | 181.2219 | 214.5277 | 8.12 | 0.003105 | 1685.1364 | 0.086589 | 122.0039 | 738.918 | 191.2934 | 124.0282 | 6.399 | 151.5257 | 179.3738 | 6.7894 | 0.0025962
739.57 | 231.1693 | 148.5525 | 7.5399 | 182.5904 | 215.5757 | 7.9816 | 0.000441 | 1702.3653 | 0.086624 | 122.1822 | 739.57 | 191.5336 | 123.0821 | 6.2472 | 151.284 | 178.6137 | 6.6131 | 0.00036539
740.222 | 233.4028 | 148.63 | 7.4162 | 183.7292 | 216.5243 | 7.8703 | -0.001096 | 1718.3944 | 0.086612 | 122.3609 | 740.222 | 191.8875 | 122.1932 | 6.0971 | 151.0494 | 178.0113 | 6.4704 | -0.00090106
740.874 | 235.0792 | 148.3219 | 7.2618 | 184.532 | 216.9299 | 7.765 | 0.004402 | 1731.0675 | 0.08664 | 122.5398 | 740.874 | 192.0693 | 121.1851 | 5.9332 | 150.7702 | 177.2406 | 6.3443 | 0.0035966
741.526 | 236.2986 | 147.6842 | 7.0811 | 184.8766 | 216.8598 | 7.6421 | 0.002416 | 1739.0693 | 0.086686 | 122.7184 | 741.526 | 192.3561 | 120.2206 | 5.7643 | 150.4966 | 176.5322 | 6.2209 | 0.0019667
742.178 | 236.8562 | 146.5361 | 6.8671 | 184.6743 | 216.1048 | 7.4986 | 0.004922 | 1742.7105 | 0.086623 | 122.8964 | 742.178 | 192.8253 | 119.2955 | 5.5905 | 150.3439 | 175.9315 | 6.1046 | 0.004007
742.83 | 236.7701 | 145.0491 | 6.6351 | 183.9799 | 214.8501 | 7.3367 | 0.003236 | 1741.0595 | 0.086652 | 123.0732 | 742.83 | 193.1517 | 118.3278 | 5.4128 | 150.0866 | 175.2699 | 5.9851 | 0.0026399
743.482 | 236.1927 | 143.27 | 6.3816 | 182.8833 | 212.9851 | 7.1753 | -0.001128 | 1736.0902 | 0.086514 | 123.2483 | 743.482 | 193.8157 | 117.5649 | 5.2367 | 150.0709 | 174.772 | 5.8879 | -0.00092562
744.134 | 235.2815 | 141.1641 | 6.1184 | 181.4536 | 210.951 | 6.9579 | 0.003846 | 1728.0673 | 0.086511 | 123.4213 | 744.134 | 194.2438 | 116.5423 | 5.0513 | 149.8045 | 174.157 | 5.7443 | 0.0031752
744.786 | 233.9752 | 138.9921 | 5.839 | 179.8195 | 208.6403 | 6.8004 | 0.002621 | 1719.4547 | 0.086594 | 123.5916 | 744.786 | 194.2141 | 115.3722 | 4.8468 | 149.2615 | 173.1846 | 5.6447 | 0.0021756
745.438 | 232.804 | 136.7039 | 5.569 | 178.1897 | 206.3691 | 6.6253 | 0.000168 | 1710.0858 | 0.086552 | 123.7587 | 745.438 | 194.6586 | 114.3047 | 4.6565 | 148.993 | 172.5551 | 5.5397 | 0.00014047
746.09 | 231.6094 | 134.4897 | 5.291 | 176.5875 | 204.0134 | 6.4264 | -0.001273 | 1700.5858 | 0.086576 | 123.9223 | 746.09 | 194.9432 | 113.1985 | 4.4534 | 148.6318 | 171.7159 | 5.409 | -0.0010715
746.742 | 230.7202 | 132.5036 | 5.0199 | 175.1854 | 202.089 | 6.2742 | 0.004234 | 1693.5255 | 0.086605 | 124.0818 | 746.742 | 195.19 | 112.0984 | 4.2469 | 148.2074 | 170.9679 | 5.308 | 0.003582
747.394 | 229.802 | 130.4194 | 4.7557 | 173.7674 | 199.9223 | 6.1487 | 0.003885 | 1686.0136 | 0.086496 | 124.2366 | 747.394 | 195.7694 | 111.1049 | 4.0514 | 148.0333 | 170.3148 | 5.2381 | 0.0033096
748.046 | 229.1849 | 128.5696 | 4.5032 | 172.6113 | 198.1746 | 6.0209 | 0.004293 | 1680.6832 | 0.086476 | 124.3864 | 748.046 | 196.1445 | 110.0344 | 3.854 | 147.7268 | 169.6048 | 5.1529 | 0.0036741
748.698 | 228.8131 | 126.7536 | 4.2573 | 171.5217 | 196.575 | 5.8756 | 0.002771 | 1677.0473 | 0.086519 | 124.5306 | 748.698 | 196.3813 | 108.7876 | 3.6539 | 147.2103 | 168.7126 | 5.0428 | 0.0023782
749.35 | 228.5562 | 125.1222 | 4.0256 | 170.6179 | 195.1133 | 5.7798 | -0.000753 | 1674.0959 | 0.086465 | 124.6688 | 749.35 | 196.848 | 107.7637 | 3.4671 | 146.9476 | 168.0447 | 4.9779 | -0.00064853
750.002 | 228.5235 | 123.4963 | 3.8 | 169.7534 | 193.7435 | 5.6836 | -0.001482 | 1672.869 | 0.086569 | 124.8004 | 750.002 | 196.9357 | 106.426 | 3.2748 | 146.2891 | 166.9631 | 4.898 | -0.0012771
750.654 | 228.5988 | 121.9688 | 3.5857 | 169.0686 | 192.5895 | 5.5447 | -0.002758 | 1673.0766 | 0.086431 | 124.9253 | 750.654 | 197.4878 | 105.3695 | 3.0977 | 146.0593 | 166.3792 | 4.7901 | -0.0023827
751.306 | 228.765 | 120.44 | 3.3757 | 168.3768 | 191.3172 | 5.437 | 0.005437 | 1672.9729 | 0.086443 | 125.0442 | 751.306 | 197.8027 | 104.1391 | 2.9189 | 145.5878 | 165.4234 | 4.7011 | 0.0047011
751.958 | 229.0951 | 119.0746 | 3.1759 | 167.8128 | 190.3718 | 5.3557 | 0.011343 | 1674.5977 | 0.086434 | 125.1583 | 751.958 | 198.0976 | 102.9633 | 2.7461 | 145.1071 | 164.6138 | 4.6311 | 0.0098082
752.61 | 229.3909 | 117.6115 | 2.9796 | 167.1931 | 189.3379 | 5.2692 | 0.003864 | 1675.3695 | 0.086426 | 125.2687 | 752.61 | 198.4552 | 101.7504 | 2.5778 | 144.6454 | 163.8037 | 4.5586 | 0.0033429
753.262 | 229.8655 | 116.2023 | 2.7974 | 166.6722 | 188.3366 | 5.1792 | 0.004089 | 1678.1384 | 0.086453 | 125.3765 | 753.262 | 198.647 | 100.4206 | 2.4175 | 144.0361 | 162.7583 | 4.4758 | 0.0035337
753.914 | 230.3492 | 114.8174 | 2.6219 | 166.1634 | 187.4465 | 5.0982 | 0.005211 | 1680.4009 | 0.086407 | 125.4827 | 753.914 | 199.0722 | 99.2274 | 2.2659 | 143.6016 | 161.9948 | 4.4059 | 0.0045034
754.566 | 230.8961 | 113.4362 | 2.4554 | 165.6035 | 186.575 | 5.0132 | 0.001545 | 1683.2389 | 0.086396 | 125.5885 | 754.566 | 199.4014 | 97.9633 | 2.1204 | 143.0149 | 161.1258 | 4.3294 | 0.0013343
755.218 | 231.4102 | 112.0876 | 2.294 | 165.1754 | 185.7226 | 4.952 | 0.004197 | 1687.0442 | 0.086374 | 125.6949 | 755.218 | 199.613 | 96.6861 | 1.9788 | 142.4793 | 160.2032 | 4.2716 | 0.0036203
755.87 | 232.09 | 110.683 | 2.145 | 164.6726 | 184.8274 | 4.8866 | 0.002749 | 1690.0988 | 0.086367 | 125.8031 | 755.87 | 200.0259 | 95.3917 | 1.8486 | 141.9224 | 159.2928 | 4.2115 | 0.0023692
756.522 | 232.6715 | 109.3019 | 2.0016 | 164.1237 | 183.929 | 4.8025 | 0.001984 | 1693.7239 | 0.08641 | 125.9142 | 756.522 | 200.1751 | 94.0361 | 1.722 | 141.2011 | 158.2402 | 4.1318 | 0.0017069
757.174 | 233.3072 | 107.9481 | 1.8628 | 163.6687 | 183.138 | 4.7393 | 0.003728 | 1697.5134 | 0.086435 | 126.0292 | 757.174 | 200.4 | 92.7224 | 1.6001 | 140.5838 | 157.307 | 4.0708 | 0.0032022
757.826 | 233.8776 | 106.5081 | 1.7399 | 163.1021 | 182.2555 | 4.6471 | 0.009096 | 1701.6226 | 0.08638 | 126.1492 | 757.826 | 200.7236 | 91.4097 | 1.4932 | 139.981 | 156.4192 | 3.9883 | 0.0078066
758.478 | 234.6036 | 105.1723 | 1.621 | 162.6544 | 181.3936 | 4.6164 | 0.002427 | 1706.1893 | 0.086451 | 126.2754 | 758.478 | 200.8437 | 90.0378 | 1.3877 | 139.2481 | 155.2907 | 3.9521 | 0.0020777
759.13 | 235.2686 | 103.7008 | 1.5093 | 162.0768 | 180.5183 | 4.5206 | 0.005465 | 1709.9558 | 0.08634 | 126.4088 | 759.13 | 201.4388 | 88.7895 | 1.2922 | 138.7714 | 154.5612 | 3.8706 | 0.0046792
759.782 | 235.8669 | 102.3063 | 1.404 | 161.506 | 179.6046 | 4.479 | 0.001896 | 1714.2987 | 0.086369 | 126.5506 | 759.782 | 201.598 | 87.4423 | 1.2001 | 138.041 | 153.51 | 3.8282 | 0.0016205
760.434 | 236.5437 | 100.791 | 1.3112 | 160.8567 | 178.6673 | 4.4198 | 0.002182 | 1717.3965 | 0.086328 | 126.7016 | 760.434 | 202.1498 | 86.1358 | 1.1206 | 137.4679 | 152.6888 | 3.7772 | 0.0018647
761.086 | 237.2163 | 99.4443 | 1.2231 | 160.3359 | 177.8197 | 4.3662 | 0.004293 | 1722.3651 | 0.086364 | 126.8616 | 761.086 | 202.3101 | 84.8111 | 1.0431 | 136.7426 | 151.6537 | 3.7237 | 0.0036613
761.738 | 237.8947 | 97.9454 | 1.1465 | 159.7145 | 176.8704 | 4.2949 | 0.006406 | 1726.7581 | 0.086451 | 127.0294 | 761.738 | 202.435 | 83.346 | 0.9756 | 135.908 | 150.5068 | 3.6548 | 0.0054511
762.39 | 238.5698 | 96.5073 | 1.0723 | 159.1104 | 175.9632 | 4.2417 | 0.000334 | 1730.7018 | 0.086349 | 127.2037 | 762.39 | 203.0653 | 82.1449 | 0.91273 | 135.4312 | 149.7759 | 3.6104 | 0.00028429
763.042 | 239.2491 | 95.0514 | 1.0072 | 158.4339 | 175.0322 | 4.1608 | 0.008719 | 1734.935 | 0.086362 | 127.3832 | 763.042 | 203.4022 | 80.8098 | 0.85633 | 134.6957 | 148.807 | 3.5373 | 0.0074126
763.6912 | 239.8313 | 93.5496 | 0.94959 | 157.7603 | 173.9751 | 4.125 | 0.003674 | 1739.1683 | 0.086348 | 127.5659 | 763.6912 | 203.7252 | 79.4659 | 0.80663 | 134.0099 | 147.7836 | 3.504 | 0.0031209
764.3403 | 240.5124 | 92.0408 | 0.89602 | 157.0044 | 172.9985 | 4.0761 | -0.001167 | 1742.7871 | 0.086307 | 127.7514 | 764.3403 | 204.2729 | 78.1725 | 0.76101 | 133.3476 | 146.9317 | 3.462 | -0.00099116
764.9895 | 241.2411 | 90.5648 | 0.85129 | 156.3412 | 172.0907 | 4.0222 | -0.001451 | 1746.6686 | 0.086276 | 127.9384 | 764.9895 | 204.8103 | 76.8883 | 0.72273 | 132.7315 | 146.1026 | 3.4148 | -0.0012319
765.6387 | 241.8221 | 89.0543 | 0.80926 | 155.5692 | 171.0199 | 3.9786 | 0.005545 | 1750.5448 | 0.086315 | 128.1257 | 765.6387 | 205.0552 | 75.5143 | 0.68622 | 131.9162 | 145.0178 | 3.3737 | 0.0047019
766.2878 | 242.5894 | 87.5654 | 0.77213 | 154.8395 | 170.0694 | 3.9181 | 0.001432 | 1755.6554 | 0.086282 | 128.3121 | 766.2878 | 205.4853 | 74.1722 | 0.65403 | 131.1568 | 144.0573 | 3.3188 | 0.001213
766.937 | 243.15 | 86.0686 | 0.74049 | 154.0641 | 169.0106 | 3.8695 | 0.011341 | 1759.9696 | 0.086239 | 128.4964 | 766.937 | 205.8524 | 72.8663 | 0.62691 | 130.4317 | 143.0855 | 3.276 | 0.0096014
767.5862 | 243.6781 | 84.5322 | 0.7113 | 153.1917 | 167.8824 | 3.818 | 0.005323 | 1764.3614 | 0.086228 | 128.6773 | 767.5862 | 206.1022 | 71.497 | 0.60162 | 129.569 | 141.9944 | 3.2293 | 0.0045022
768.2354 | 244.2892 | 82.9738 | 0.68586 | 152.3463 | 166.7578 | 3.752 | 0.000183 | 1768.1257 | 0.086309 | 128.8536 | 768.2354 | 206.2671 | 70.0594 | 0.57911 | 128.6345 | 140.803 | 3.168 | 0.00015452
768.8845 | 244.8441 | 81.4864 | 0.66313 | 151.4482 | 165.6295 | 3.695 | 0.004571 | 1772.9057 | 0.086279 | 129.0241 | 768.8845 | 206.5236 | 68.733 | 0.55935 | 127.7451 | 139.7069 | 3.1167 | 0.0038556
769.5337 | 245.5032 | 79.9816 | 0.64365 | 150.6155 | 164.5445 | 3.663 | 0.007625 | 1777.5073 | 0.086267 | 129.1876 | 769.5337 | 206.8345 | 67.3839 | 0.54227 | 126.8924 | 138.6274 | 3.0861 | 0.006424
770.1829 | 246.1886 | 78.5235 | 0.62695 | 149.7715 | 163.5323 | 3.6021 | -0.00178 | 1782.7412 | 0.08629 | 129.3428 | 770.1829 | 206.9955 | 66.0226 | 0.52714 | 125.9279 | 137.498 | 3.0286 | -0.0014966
770.832 | 246.8082 | 77.0114 | 0.61206 | 148.8311 | 162.3961 | 3.5304 | 0.013723 | 1785.7957 | 0.086267 | 129.4893 | 770.832 | 207.4524 | 64.7313 | 0.51446 | 125.0986 | 136.5005 | 2.9675 | 0.011535
771.4811 | 247.4661 | 75.534 | 0.5981 | 148.0071 | 161.385 | 3.5071 | 0.007833 | 1788.9232 | 0.086222 | 129.6277 | 771.4811 | 207.9714 | 63.479 | 0.50265 | 124.3857 | 135.6285 | 2.9474 | 0.0065829
772.1303 | 248.1442 | 74.0324 | 0.58567 | 147.0703 | 160.2681 | 3.4496 | 0.007521 | 1791.5637 | 0.086213 | 129.759 | 772.1303 | 208.4663 | 62.1947 | 0.49202 | 123.554 | 134.6414 | 2.898 | 0.0063184
772.7795 | 248.8889 | 72.5614 | 0.57623 | 146.17 | 159.2407 | 3.4404 | 0.007141 | 1796.3354 | 0.086269 | 129.8839 | 772.7795 | 208.6017 | 60.816 | 0.48295 | 122.5098 | 133.4647 | 2.8835 | 0.0059851
773.4286 | 249.5647 | 71.102 | 0.56422 | 145.2233 | 158.1121 | 3.3834 | 0.006035 | 1800.2455 | 0.086199 | 130.0034 | 773.4286 | 209.0764 | 59.5667 | 0.47268 | 121.6629 | 132.4607 | 2.8345 | 0.0050559
774.0778 | 250.1352 | 69.6293 | 0.55693 | 144.2621 | 156.9559 | 3.3272 | 0.008294 | 1804.8423 | 0.086249 | 130.1183 | 774.0778 | 209.0832 | 58.2018 | 0.46553 | 120.5859 | 131.1964 | 2.7811 | 0.0069328
774.727 | 250.8266 | 68.1684 | 0.54904 | 143.2519 | 155.8361 | 3.2768 | 0.007344 | 1809.6736 | 0.086322 | 130.2294 | 774.727 | 209.1025 | 56.8289 | 0.4577 | 119.4225 | 129.9133 | 2.7317 | 0.0061224
775.3762 | 251.4938 | 66.6912 | 0.53978 | 142.2418 | 154.6765 | 3.2415 | 0.013144 | 1813.7648 | 0.086236 | 130.3376 | 775.3762 | 209.5697 | 55.5737 | 0.4498 | 118.53 | 128.8918 | 2.7011 | 0.010953
776.0253 | 252.1463 | 65.2488 | 0.5327 | 141.2501 | 153.5156 | 3.1895 | 0.008678 | 1818.2822 | 0.086277 | 130.4438 | 776.0253 | 209.662 | 54.255 | 0.44295 | 117.4508 | 127.6497 | 2.6521 | 0.0072158
776.6745 | 252.7464 | 63.7865 | 0.52488 | 140.1834 | 152.3049 | 3.1424 | 0.014582 | 1822.4597 | 0.086285 | 130.5488 | 776.6745 | 209.8292 | 52.9553 | 0.43576 | 116.3798 | 126.443 | 2.6088 | 0.012106
777.3237 | 253.4264 | 62.3408 | 0.51682 | 139.1618 | 151.113 | 3.1041 | 0.014752 | 1827.3294 | 0.086301 | 130.6535 | 777.3237 | 209.962 | 51.6489 | 0.42818 | 115.2946 | 125.196 | 2.5717 | 0.012222
777.9728 | 253.9984 | 60.934 | 0.50865 | 138.0827 | 149.9331 | 3.0876 | 0.011658 | 1831.9446 | 0.086188 | 130.7587 | 777.9728 | 210.3491 | 50.4626 | 0.42124 | 114.3534 | 124.1672 | 2.557 | 0.0096546
778.622 | 254.8287 | 59.5589 | 0.50035 | 137.0802 | 148.7824 | 3.05 | 0.00889 | 1837.7092 | 0.086291 | 130.8653 | 778.622 | 210.2951 | 49.1505 | 0.41291 | 113.1242 | 122.7813 | 2.517 | 0.0073364
779.2712 | 255.5648 | 58.1708 | 0.49271 | 136.023 | 147.6325 | 2.9844 | 0.008031 | 1842.9485 | 0.086335 | 130.9741 | 779.2712 | 210.3719 | 47.8842 | 0.40558 | 111.9693 | 121.5259 | 2.4566 | 0.0066108
779.9203 | 256.1406 | 56.7659 | 0.48578 | 134.8837 | 146.4311 | 2.9502 | 0.011149 | 1846.5991 | 0.086274 | 131.086 | 779.9203 | 210.7574 | 46.7081 | 0.39971 | 110.985 | 120.4864 | 2.4275 | 0.0091736
780.5695 | 256.919 | 55.3861 | 0.4768 | 133.8242 | 145.2306 | 2.9241 | 0.017112 | 1851.0028 | 0.086273 | 131.2017 | 780.5695 | 211.0833 | 45.5049 | 0.39174 | 109.9492 | 119.3207 | 2.4024 | 0.014059
781.2187 | 257.3671 | 53.9968 | 0.46875 | 132.5999 | 143.9514 | 2.8993 | 0.014153 | 1854.7777 | 0.0863 | 131.321 | 781.2187 | 211.1476 | 44.2998 | 0.38457 | 108.7868 | 118.0998 | 2.3787 | 0.011611
781.8679 | 258.1254 | 52.6808 | 0.45893 | 131.5362 | 142.7823 | 2.8561 | 0.014748 | 1859.6713 | 0.086288 | 131.4433 | 781.8679 | 211.4387 | 43.1525 | 0.37592 | 107.7455 | 116.9575 | 2.3395 | 0.012081
782.517 | 258.627 | 51.3041 | 0.45096 | 130.2718 | 141.4583 | 2.8191 | 0.015017 | 1862.6653 | 0.086241 | 131.5681 | 782.517 | 211.8236 | 42.0197 | 0.36935 | 106.6967 | 115.8588 | 2.3089 | 0.012299
783.1661 | 259.2709 | 50.0054 | 0.44174 | 129.1085 | 140.2461 | 2.7833 | 0.010861 | 1866.8876 | 0.086282 | 131.6947 | 783.1661 | 211.9737 | 40.8832 | 0.36116 | 105.556 | 114.6618 | 2.2756 | 0.0088797
783.8153 | 259.9269 | 48.6878 | 0.43523 | 127.8589 | 138.9519 | 2.733 | 0.010853 | 1870.6595 | 0.086224 | 131.8227 | 783.8153 | 212.4306 | 39.7911 | 0.3557 | 104.4953 | 113.5613 | 2.2336 | 0.0088698
784.4645 | 260.5371 | 47.4433 | 0.42494 | 126.6869 | 137.7642 | 2.6888 | 0.016937 | 1874.4379 | 0.086297 | 131.9514 | 784.4645 | 212.5294 | 38.7012 | 0.34664 | 103.343 | 112.3792 | 2.1934 | 0.013816
785.1136 | 261.1028 | 46.1251 | 0.41746 | 125.4645 | 136.5166 | 2.6835 | 0.015854 | 1878.2714 | 0.086297 | 132.0803 | 785.1136 | 212.7631 | 37.5857 | 0.34018 | 102.2364 | 111.2424 | 2.1867 | 0.012919
785.7628 | 261.737 | 44.8728 | 0.40924 | 124.2057 | 135.267 | 2.6485 | 0.020645 | 1882.5736 | 0.086268 | 132.2089 | 785.7628 | 213.0698 | 36.5292 | 0.33315 | 101.111 | 110.1156 | 2.156 | 0.016806
786.412 | 262.2659 | 43.6327 | 0.40127 | 122.9366 | 133.9881 | 2.5897 | 0.009214 | 1886.7667 | 0.08628 | 132.3365 | 786.412 | 213.202 | 35.47 | 0.3262 | 99.938 | 108.922 | 2.1053 | 0.0074903
787.0612 | 262.9112 | 42.4225 | 0.39169 | 121.6878 | 132.7609 | 2.5743 | 0.0173 | 1891.5428 | 0.086255 | 132.4627 | 787.0612 | 213.4542 | 34.4422 | 0.318 | 98.7967 | 107.7868 | 2.09 | 0.014046
787.7103 | 263.5187 | 41.2049 | 0.3847 | 120.4209 | 131.5161 | 2.5372 | 0.010373 | 1894.9555 | 0.086194 | 132.5868 | 787.7103 | 213.9126 | 33.4482 | 0.31228 | 97.7522 | 106.7588 | 2.0596 | 0.0084203
788.3595 | 264.0582 | 40.0489 | 0.37703 | 119.168 | 130.3075 | 2.5283 | 0.013333 | 1898.981 | 0.086275 | 132.7083 | 788.3595 | 213.891 | 32.4402 | 0.3054 | 96.5278 | 105.551 | 2.048 | 0.0108
789.0087 | 264.5909 | 38.8575 | 0.36761 | 117.8063 | 128.9493 | 2.4732 | 0.010694 | 1903.2313 | 0.08622 | 132.8267 | 789.0087 | 214.1713 | 31.4529 | 0.29756 | 95.3575 | 104.3771 | 2.0019 | 0.0086562
789.6578 | 265.2136 | 37.7226 | 0.36076 | 116.524 | 127.6925 | 2.4568 | 0.012975 | 1906.8026 | 0.086238 | 132.9413 | 789.6578 | 214.4139 | 30.4971 | 0.29166 | 94.2047 | 103.2339 | 1.9863 | 0.01049
790.307 | 265.8262 | 36.5822 | 0.35135 | 115.2189 | 126.4672 | 2.4058 | 0.008024 | 1910.376 | 0.086265 | 133.0516 | 790.307 | 214.6171 | 29.535 | 0.28367 | 93.023 | 102.1044 | 1.9423 | 0.0064782
790.9562 | 266.3273 | 35.467 | 0.34428 | 113.8811 | 125.2156 | 2.38 | 0.014283 | 1912.7589 | 0.086246 | 133.1575 | 790.9562 | 214.9722 | 28.628 | 0.27789 | 91.9218 | 101.0707 | 1.9211 | 0.011529
791.6053 | 266.8468 | 34.3632 | 0.33597 | 112.5147 | 123.9055 | 2.3536 | 0.007049 | 1915.2921 | 0.0863 | 133.2592 | 791.6053 | 215.1362 | 27.7041 | 0.27086 | 90.7112 | 99.8946 | 1.8975 | 0.005683
792.2545 | 267.4879 | 33.2943 | 0.32744 | 111.2155 | 122.5977 | 2.3312 | 0.004581 | 1919.4415 | 0.086262 | 133.3568 | 792.2545 | 215.4399 | 26.8158 | 0.26373 | 89.5751 | 98.7425 | 1.8776 | 0.0036896
792.9037 | 268.0287 | 32.2546 | 0.3207 | 109.9334 | 121.3846 | 2.3015 | 0.003355 | 1923.8902 | 0.086202 | 133.4508 | 792.9037 | 215.6777 | 25.9547 | 0.25806 | 88.4614 | 97.6759 | 1.8519 | 0.0026997
793.5529 | 268.5879 | 31.2155 | 0.31246 | 108.5681 | 120.1166 | 2.2671 | 0.00038 | 1927.2932 | 0.086251 | 133.5412 | 793.5529 | 215.7705 | 25.077 | 0.25101 | 87.2183 | 96.4959 | 1.8213 | 0.00030527
794.202 | 269.2203 | 30.2089 | 0.30479 | 107.2139 | 118.8525 | 2.2327 | 0.000969 | 1930.836 | 0.086212 | 133.6283 | 794.202 | 216.1184 | 24.2504 | 0.24467 | 86.0667 | 95.4096 | 1.7923 | 0.00077787
794.848 | 269.7439 | 29.1979 | 0.29733 | 105.8688 | 117.5535 | 2.2068 | -0.003962 | 1933.7162 | 0.08624 | 133.712 | 794.848 | 216.2809 | 23.4109 | 0.2384 | 84.8857 | 94.2545 | 1.7694 | -0.0031767
795.494 | 270.4005 | 28.2385 | 0.29114 | 104.5473 | 116.3311 | 2.1843 | 0.007792 | 1937.7152 | 0.08627 | 133.7929 | 795.494 | 216.4175 | 22.6009 | 0.23302 | 83.6754 | 93.1067 | 1.7482 | 0.0062364
796.14 | 270.9177 | 27.2878 | 0.28371 | 103.1974 | 115.0648 | 2.1527 | 0.008987 | 1941.5459 | 0.08626 | 133.8713 | 796.14 | 216.5539 | 21.8121 | 0.22678 | 82.4893 | 91.9753 | 1.7207 | 0.0071836
796.786 | 271.4573 | 26.3626 | 0.27563 | 101.842 | 113.7983 | 2.12 | 0.016484 | 1944.9195 | 0.086246 | 133.9474 | 796.786 | 216.7672 | 21.0514 | 0.2201 | 81.3241 | 90.8715 | 1.6929 | 0.013163
797.432 | 272.1131 | 25.4794 | 0.26927 | 100.5305 | 112.5529 | 2.085 | 0.015705 | 1949.2937 | 0.086312 | 134.0214 | 797.432 | 216.7581 | 20.2962 | 0.21449 | 80.08 | 89.6567 | 1.6608 | 0.01251
798.078 | 272.6712 | 24.5734 | 0.26194 | 99.1782 | 111.3014 | 2.0664 | 0.023313 | 1952.8055 | 0.086216 | 134.0936 | 798.078 | 217.1697 | 19.5715 | 0.20862 | 78.9908 | 88.6463 | 1.6458 | 0.018568
798.724 | 273.1572 | 23.7152 | 0.25672 | 97.8112 | 110.0543 | 2.0478 | 0.039105 | 1957.1339 | 0.086224 | 134.1643 | 798.724 | 217.1701 | 18.8545 | 0.20411 | 77.7635 | 87.4973 | 1.6281 | 0.03109
799.37 | 273.8144 | 22.8777 | 0.24884 | 96.4463 | 108.8854 | 2.0167 | 0.062223 | 1960.7345 | 0.086287 | 134.2335 | 799.37 | 217.2471 | 18.1514 | 0.19743 | 76.5214 | 86.3908 | 1.6 | 0.049368
800.016 | 274.2412 | 22.062 | 0.24141 | 95.0795 | 107.5811 | 2.0207 | 0.14632 | 1963.451 | 0.086287 | 134.3017 | 800.016 | 217.3957 | 17.4889 | 0.19137 | 75.3711 | 85.2814 | 1.6018 | 0.11599
800.662 | 274.7341 | 21.2466 | 0.23474 | 93.6759 | 106.2866 | 1.9674 | 0.30025 | 1966.7409 | 0.086335 | 134.3689 | 800.662 | 217.4082 | 16.8133 | 0.18576 | 74.1295 | 84.1089 | 1.5569 | 0.2376
801.308 | 275.3412 | 20.4491 | 0.22952 | 92.3342 | 105.0948 | 1.947 | 0.49738 | 1970.0965 | 0.086296 | 134.4356 | 801.308 | 217.724 | 16.17 | 0.18149 | 73.0125 | 83.1029 | 1.5396 | 0.3933
801.954 | 275.7805 | 19.6905 | 0.22053 | 90.9936 | 103.8244 | 1.9237 | 0.71812 | 1973.7351 | 0.086318 | 134.5019 | 801.954 | 217.7204 | 15.5451 | 0.1741 | 71.8367 | 81.9663 | 1.5187 | 0.56693
802.6 | 276.3095 | 18.96 | 0.21507 | 89.5685 | 102.6013 | 1.8953 | 0.91745 | 1977.1162 | 0.086227 | 134.568 | 802.6 | 218.1021 | 14.9659 | 0.16976 | 70.7 | 80.9873 | 1.496 | 0.72418
803.246 | 276.7805 | 18.235 | 0.20939 | 88.2188 | 101.4155 | 1.8788 | 1.0781 | 1979.0697 | 0.086245 | 134.6343 | 803.246 | 218.3218 | 14.3836 | 0.16516 | 69.5862 | 79.9956 | 1.482 | 0.85038
803.892 | 277.4208 | 17.5346 | 0.20172 | 86.87 | 100.196 | 1.8446 | 1.2331 | 1982.8335 | 0.086268 | 134.7011 | 803.892 | 218.4623 | 13.8081 | 0.15885 | 68.4081 | 78.902 | 1.4526 | 0.97103
804.538 | 277.8575 | 16.838 | 0.19838 | 85.4957 | 98.9767 | 1.8128 | 1.4788 | 1985.9142 | 0.086296 | 134.7684 | 804.538 | 218.5046 | 13.2412 | 0.156 | 67.233 | 77.8343 | 1.4255 | 1.1629
805.184 | 278.4033 | 16.1532 | 0.19156 | 84.1582 | 97.8744 | 1.8194 | 2.0921 | 1989.7983 | 0.086285 | 134.8367 | 805.184 | 218.644 | 12.6859 | 0.15044 | 66.0937 | 76.8657 | 1.4289 | 1.643
805.83 | 278.8286 | 15.5079 | 0.18674 | 82.7701 | 96.6043 | 1.7728 | 6.1584 | 1992.3067 | 0.086306 | 134.9062 | 805.83 | 218.7611 | 12.1671 | 0.14651 | 64.9391 | 75.793 | 1.3909 | 4.8317
806.476 | 279.2959 | 14.8789 | 0.18173 | 81.4071 | 95.473 | 1.7596 | 44.3617 | 1995.877 | 0.08638 | 134.977 | 806.476 | 218.6634 | 11.6488 | 0.14228 | 63.7344 | 74.7467 | 1.3776 | 34.7312
807.122 | 279.7524 | 14.2504 | 0.17717 | 80.0814 | 94.2294 | 1.7316 | 233.0438 | 1998.9946 | 0.086314 | 135.0496 | 807.122 | 218.9651 | 11.1539 | 0.13867 | 62.6806 | 73.7543 | 1.3554 | 182.4058
807.768 | 280.2545 | 13.6681 | 0.17111 | 78.7063 | 93.0766 | 1.7025 | 491.6758 | 2001.6995 | 0.086304 | 135.1241 | 807.768 | 219.2073 | 10.6908 | 0.13384 | 61.5619 | 72.8019 | 1.3317 | 384.5751
808.414 | 280.7335 | 13.0991 | 0.16599 | 77.3449 | 91.8898 | 1.67 | 616.0431 | 2004.1618 | 0.086306 | 135.2008 | 808.414 | 219.4306 | 10.2387 | 0.12975 | 60.4554 | 71.8242 | 1.3053 | 481.5198
809.06 | 281.0175 | 12.5663 | 0.16037 | 76.0071 | 90.6803 | 1.6556 | 631.9073 | 2006.715 | 0.086275 | 135.2799 | 809.06 | 219.5823 | 9.8191 | 0.12531 | 59.3906 | 70.856 | 1.2937 | 493.7616
809.706 | 281.4925 | 12.0204 | 0.1561 | 74.6592 | 89.5471 | 1.6327 | 476.1973 | 2009.5254 | 0.086262 | 135.3618 | 809.706 | 219.8114 | 9.3865 | 0.1219 | 58.2998 | 69.9254 | 1.2749 | 371.8522
810.352 | 281.9834 | 11.5191 | 0.15274 | 73.3424 | 88.4409 | 1.6245 | 211.7125 | 2012.8698 | 0.086298 | 135.4466 | 810.352 | 219.8747 | 8.982 | 0.1191 | 57.1882 | 68.9612 | 1.2667 | 165.0814
810.998 | 282.4325 | 11.0245 | 0.14791 | 72.0243 | 87.2453 | 1.6055 | 70.0502 | 2016.0601 | 0.086275 | 135.5342 | 810.998 | 220.0784 | 8.5905 | 0.11525 | 56.1231 | 67.9837 | 1.251 | 54.5849
811.644 | 282.7334 | 10.5484 | 0.14393 | 70.7168 | 86.0871 | 1.5774 | 23.472 | 2017.3594 | 0.086296 | 135.6242 | 811.644 | 220.2631 | 8.2177 | 0.11213 | 55.0919 | 67.066 | 1.2289 | 18.2859
812.29 | 283.2351 | 10.0786 | 0.13917 | 69.4283 | 84.9785 | 1.5635 | 8.0538 | 2020.6984 | 0.086366 | 135.7161 | 812.29 | 220.2601 | 7.8377 | 0.10823 | 53.9915 | 66.0843 | 1.2158 | 6.2631
812.936 | 283.6105 | 9.6284 | 0.13657 | 68.1181 | 83.835 | 1.5365 | 2.7315 | 2022.9579 | 0.086319 | 135.8095 | 812.936 | 220.5754 | 7.4884 | 0.10621 | 52.9782 | 65.2019 | 1.195 | 2.1244
813.582 | 284.0066 | 9.2012 | 0.13306 | 66.8337 | 82.7125 | 1.501 | 1.3182 | 2025.2953 | 0.086364 | 135.904 | 813.582 | 220.6676 | 7.1492 | 0.10338 | 51.9285 | 64.266 | 1.1662 | 1.0242
814.228 | 284.26 | 8.7832 | 0.12894 | 65.5509 | 81.5778 | 1.4949 | 0.92357 | 2028.0906 | 0.086295 | 135.999 | 814.228 | 220.8916 | 6.8252 | 0.1002 | 50.9381 | 63.3921 | 1.1617 | 0.71768
814.874 | 284.723 | 8.4011 | 0.12704 | 64.2755 | 80.4905 | 1.4702 | 0.73376 | 2030.2549 | 0.086279 | 136.0941 | 814.874 | 221.2113 | 6.5271 | 0.098699 | 49.9379 | 62.5359 | 1.1422 | 0.57008
815.52 | 285.0593 | 8.0228 | 0.1233 | 63.0243 | 79.3811 | 1.4403 | 0.54696 | 2032.5964 | 0.086402 | 136.1889 | 815.52 | 221.0559 | 6.2215 | 0.095612 | 48.8737 | 61.5579 | 1.1169 | 0.42415
816.166 | 285.4665 | 7.6642 | 0.12048 | 61.7916 | 78.3113 | 1.4393 | 0.38382 | 2035.1067 | 0.086244 | 136.2829 | 816.166 | 221.6569 | 5.951 | 0.093549 | 47.9795 | 60.8066 | 1.1176 | 0.29802
816.812 | 285.8385 | 7.3115 | 0.11695 | 60.5205 | 77.2495 | 1.4174 | 0.24901 | 2037.1365 | 0.086315 | 136.3756 | 816.812 | 221.6914 | 5.6707 | 0.090706 | 46.9387 | 59.9133 | 1.0993 | 0.19313
817.458 | 286.1928 | 6.9735 | 0.11483 | 59.3072 | 76.1115 | 1.3716 | 0.15609 | 2039.3947 | 0.08629 | 136.4667 | 817.458 | 221.9337 | 5.4077 | 0.08905 | 45.9909 | 59.0221 | 1.0636 | 0.12104
818.104 | 286.6166 | 6.6346 | 0.11137 | 58.1237 | 75.098 | 1.3637 | 0.10914 | 2042.3188 | 0.086347 | 136.5556 | 818.104 | 221.9421 | 5.1375 | 0.086241 | 45.0082 | 58.1523 | 1.056 | 0.084514
818.75 | 286.8999 | 6.3231 | 0.11055 | 56.911 | 74.0778 | 1.3501 | 0.09223 | 2043.8386 | 0.086409 | 136.642 | 818.75 | 221.9777 | 4.8922 | 0.085531 | 44.0327 | 57.3148 | 1.0446 | 0.071359
819.396 | 287.3011 | 6.0144 | 0.10732 | 55.7237 | 73.0313 | 1.3479 | 0.08884 | 2047.3715 | 0.086409 | 136.7253 | 819.396 | 222.0386 | 4.6482 | 0.08294 | 43.0657 | 56.4418 | 1.0417 | 0.068659
820.042 | 287.6118 | 5.7242 | 0.10548 | 54.5547 | 71.9779 | 1.3247 | 0.066012 | 2049.3333 | 0.08639 | 136.8051 | 820.042 | 222.245 | 4.4232 | 0.081511 | 42.1558 | 55.6192 | 1.0236 | 0.051009
820.688 | 287.9868 | 5.4441 | 0.10321 | 53.3812 | 70.9943 | 1.3083 | 0.061893 | 2051.6107 | 0.086315 | 136.8811 | 820.688 | 222.6051 | 4.2081 | 0.079778 | 41.262 | 54.8764 | 1.0113 | 0.047841
821.334 | 288.2121 | 5.1772 | 0.10153 | 52.2048 | 69.9415 | 1.2727 | 0.062683 | 2053.8849 | 0.086416 | 136.9536 | 821.334 | 222.39 | 3.9948 | 0.078339 | 40.2822 | 53.9682 | 0.98203 | 0.048367
821.98 | 288.4325 | 4.9134 | 0.098898 | 51.0596 | 68.8914 | 1.2756 | 0.059985 | 2055.5091 | 0.086398 | 137.0231 | 821.98 | 222.5424 | 3.791 | 0.076306 | 39.3954 | 53.1537 | 0.9842 | 0.046282
822.626 | 288.6942 | 4.6659 | 0.097667 | 49.9133 | 67.9544 | 1.2643 | 0.047313 | 2056.8015 | 0.086304 | 137.09 | 822.626 | 222.9563 | 3.6034 | 0.075427 | 38.5477 | 52.4807 | 0.97643 | 0.036539
823.2664 | 288.8915 | 4.434 | 0.095246 | 48.7766 | 66.8895 | 1.2547 | 0.055106 | 2057.4992 | 0.086414 | 137.1541 | 823.2664 | 222.8529 | 3.4204 | 0.073473 | 37.6266 | 51.599 | 0.96792 | 0.042509
823.9069 | 289.0697 | 4.2093 | 0.094999 | 47.6608 | 65.9267 | 1.2282 | 0.052532 | 2059.1322 | 0.086355 | 137.2164 | 823.9069 | 223.0668 | 3.2482 | 0.073308 | 36.7785 | 50.8738 | 0.94774 | 0.040537
824.5473 | 289.3533 | 3.9882 | 0.093328 | 46.5734 | 64.9696 | 1.2021 | 0.044496 | 2060.1872 | 0.086443 | 137.2774 | 824.5473 | 223.0432 | 3.0742 | 0.07194 | 35.9003 | 50.0808 | 0.92665 | 0.034299
825.1877 | 289.4358 | 3.7817 | 0.091972 | 45.4617 | 63.982 | 1.1835 | 0.038486 | 2061.0917 | 0.08639 | 137.3374 | 825.1877 | 223.2442 | 2.9169 | 0.070939 | 35.065 | 49.3498 | 0.91281 | 0.029685
825.8282 | 289.3994 | 3.5781 | 0.090076 | 44.3729 | 62.9734 | 1.1755 | 0.035572 | 2060.899 | 0.086378 | 137.397 | 825.8282 | 223.3651 | 2.7616 | 0.069523 | 34.248 | 48.6043 | 0.9073 | 0.027455
826.4686 | 289.5618 | 3.3953 | 0.089109 | 43.283 | 61.9359 | 1.1523 | 0.0395 | 2061.9354 | 0.086429 | 137.4563 | 826.4686 | 223.3417 | 2.6188 | 0.068731 | 33.3846 | 47.7718 | 0.88876 | 0.030467
827.1091 | 289.6371 | 3.2172 | 0.088188 | 42.2217 | 61.0309 | 1.1477 | 0.037187 | 2062.8449 | 0.08633 | 137.516 | 827.1091 | 223.6546 | 2.4843 | 0.068098 | 32.6031 | 47.1274 | 0.88626 | 0.028715
827.7495 | 289.8794 | 3.0485 | 0.086508 | 41.1904 | 60.1231 | 1.1256 | 0.035953 | 2064.65 | 0.086425 | 137.5764 | 827.7495 | 223.4981 | 2.3504 | 0.066698 | 31.7579 | 46.3551 | 0.86786 | 0.02772
828.3899 | 290.0714 | 2.8858 | 0.085788 | 40.1848 | 59.19 | 1.1194 | 0.029008 | 2065.8488 | 0.086357 | 137.6378 | 828.3899 | 223.7938 | 2.2264 | 0.066187 | 31.0031 | 45.6658 | 0.86362 | 0.02238
829.0303 | 290.5155 | 2.7356 | 0.084691 | 39.213 | 58.3151 | 1.1035 | 0.033538 | 2067.6964 | 0.086425 | 137.7008 | 829.0303 | 223.8618 | 2.108 | 0.06526 | 30.2163 | 44.9357 | 0.85033 | 0.025843
829.6708 | 291.0861 | 2.5883 | 0.085146 | 38.2807 | 57.5324 | 1.0842 | 0.033828 | 2071.4767 | 0.086395 | 137.7657 | 829.6708 | 224.0757 | 1.9924 | 0.065545 | 29.4681 | 44.288 | 0.83464 | 0.026041
830.3112 | 291.7605 | 2.4566 | 0.083388 | 37.3758 | 56.7302 | 1.0508 | 0.027126 | 2073.986 | 0.086366 | 137.833 | 830.3112 | 224.5075 | 1.8904 | 0.064166 | 28.7604 | 43.6534 | 0.80858 | 0.020873
830.9517 | 292.6115 | 2.3259 | 0.08259 | 36.5077 | 55.9933 | 1.061 | 0.025673 | 2079.386 | 0.086339 | 137.9026 | 830.9517 | 224.7625 | 1.7866 | 0.06344 | 28.0425 | 43.0099 | 0.81499 | 0.01972
831.5921 | 293.7151 | 2.2061 | 0.082667 | 35.6666 | 55.3141 | 1.0423 | 0.022246 | 2085.8662 | 0.086388 | 137.9745 | 831.5921 | 224.8974 | 1.6892 | 0.063298 | 27.3099 | 42.354 | 0.79809 | 0.017034
832.2325 | 294.9657 | 2.0948 | 0.081267 | 34.8878 | 54.6097 | 1.0344 | 0.026466 | 2095.141 | 0.086488 | 138.0482 | 832.2325 | 224.7154 | 1.5959 | 0.061912 | 26.5788 | 41.6037 | 0.78802 | 0.020163
832.873 | 296.2476 | 1.9822 | 0.079697 | 34.0834 | 53.9953 | 1.023 | 0.019595 | 2104.7716 | 0.086487 | 138.1236 | 832.873 | 224.7858 | 1.504 | 0.060472 | 25.8617 | 40.9704 | 0.77626 | 0.014868
833.5134 | 297.461 | 1.8778 | 0.08014 | 33.3167 | 53.3223 | 1.0145 | 0.025044 | 2114.9961 | 0.086589 | 138.2004 | 833.5134 | 224.4752 | 1.4171 | 0.060477 | 25.142 | 40.239 | 0.76561 | 0.018899
834.1538 | 298.4182 | 1.7831 | 0.080134 | 32.5309 | 52.6207 | 1.0146 | 0.019262 | 2124.3011 | 0.0867 | 138.2784 | 834.1538 | 224.0507 | 1.3387 | 0.060164 | 24.424 | 39.5073 | 0.76172 | 0.014462
834.7943 | 299.2588 | 1.6851 | 0.079607 | 31.7208 | 51.8934 | 0.99646 | 0.01584 | 2133.0191 | 0.08677 | 138.3572 | 834.7943 | 223.7096 | 1.2597 | 0.05951 | 23.7127 | 38.7927 | 0.7449 | 0.011841
835.4347 | 299.481 | 1.5963 | 0.079774 | 30.8818 | 51.1299 | 0.97273 | 0.017716 | 2141.2233 | 0.086974 | 138.4367 | 835.4347 | 222.6234 | 1.1866 | 0.059301 | 22.9564 | 38.0081 | 0.72309 | 0.013169
836.0751 | 299.4866 | 1.5071 | 0.079068 | 30.0045 | 50.2907 | 0.95357 | 0.020379 | 2146.1149 | 0.087171 | 138.5166 | 836.0751 | 221.7462 | 1.1159 | 0.058544 | 22.216 | 37.2363 | 0.70604 | 0.015089
836.7156 | 299.0474 | 1.4155 | 0.077524 | 29.1195 | 49.343 | 0.94636 | 0.019753 | 2150.505 | 0.087425 | 138.5967 | 836.7156 | 220.4535 | 1.0435 | 0.05715 | 21.4665 | 36.375 | 0.69764 | 0.014562
837.356 | 298.2809 | 1.3364 | 0.078235 | 28.2427 | 48.4252 | 0.94598 | 0.018325 | 2152.7953 | 0.08775 | 138.6766 | 837.356 | 218.9668 | 0.98102 | 0.057432 | 20.7329 | 35.5487 | 0.69444 | 0.013452
837.9965 | 297.1391 | 1.2537 | 0.077528 | 27.3297 | 47.4736 | 0.9258 | 0.024962 | 2153.5463 | 0.088046 | 138.7562 | 837.9965 | 217.4453 | 0.91747 | 0.056735 | 19.9998 | 34.741 | 0.6775 | 0.018267
838.6368 | 295.7793 | 1.1756 | 0.077324 | 26.4388 | 46.4577 | 0.89037 | 0.012472 | 2154.2826 | 0.088362 | 138.8352 | 838.6368 | 215.7234 | 0.85738 | 0.056395 | 19.2829 | 33.8834 | 0.64938 | 0.0090963
839.2773 | 294.1087 | 1.1031 | 0.0762 | 25.5218 | 45.4483 | 0.88787 | 0.009702 | 2152.9261 | 0.088513 | 138.9133 | 839.2773 | 214.3964 | 0.80416 | 0.055548 | 18.6047 | 33.1304 | 0.64723 | 0.0070725
839.9177 | 292.5113 | 1.0408 | 0.075414 | 24.6563 | 44.4878 | 0.87917 | 0.005354 | 2152.1103 | 0.088506 | 138.9902 | 839.9177 | 213.4473 | 0.75948 | 0.05503 | 17.9919 | 32.463 | 0.64153 | 0.0039068
840.5582 | 290.5691 | 0.97171 | 0.073847 | 23.7697 | 43.4058 | 0.84102 | 0.011214 | 2149.0756 | 0.088382 | 139.0658 | 840.5582 | 212.7421 | 0.71144 | 0.054068 | 17.4031 | 31.7798 | 0.61576 | 0.0082104
841.1986 | 288.7342 | 0.91224 | 0.075127 | 22.9125 | 42.3971 | 0.8063 | 0.009106 | 2149.3505 | 0.088132 | 139.14 | 841.1986 | 212.0843 | 0.67007 | 0.055183 | 16.8299 | 31.142 | 0.59225 | 0.0066886
841.8391 | 286.8223 | 0.85608 | 0.073388 | 22.0609 | 41.3472 | 0.82208 | 0.010565 | 2148.3131 | 0.087799 | 139.2127 | 841.8391 | 211.6928 | 0.63184 | 0.054165 | 16.2824 | 30.5169 | 0.60675 | 0.0077976
842.4795 | 284.8293 | 0.79876 | 0.072503 | 21.2493 | 40.3861 | 0.80219 | 0.012619 | 2147.1342 | 0.08758 | 139.2839 | 842.4795 | 210.9693 | 0.59163 | 0.053702 | 15.7391 | 29.9135 | 0.59417 | 0.0093467
843.1199 | 283.0171 | 0.74722 | 0.072671 | 20.4613 | 39.4283 | 0.77326 | 0.011173 | 2147.0153 | 0.087236 | 139.3536 | 843.1199 | 210.5729 | 0.55595 | 0.054069 | 15.2238 | 29.3358 | 0.57533 | 0.008313
843.7604 | 281.3529 | 0.69978 | 0.072551 | 19.714 | 38.5546 | 0.76081 | 0.013261 | 2146.2979 | 0.087195 | 139.4218 | 843.7604 | 209.6043 | 0.52132 | 0.05405 | 14.6867 | 28.7227 | 0.56679 | 0.0098793
844.4008 | 279.6667 | 0.66159 | 0.072543 | 18.9835 | 37.6562 | 0.74613 | 0.008692 | 2144.475 | 0.087032 | 139.4884 | 844.4008 | 209.0168 | 0.49446 | 0.054217 | 14.1878 | 28.1434 | 0.55764 | 0.0064962
845.0412 | 277.9721 | 0.62051 | 0.071281 | 18.2975 | 36.7859 | 0.73205 | 0.010387 | 2143.7544 | 0.086875 | 139.5534 | 845.0412 | 208.2912 | 0.46496 | 0.053413 | 13.7108 | 27.5646 | 0.54854 | 0.0077832
845.6816 | 276.303 | 0.58116 | 0.072363 | 17.6132 | 35.8985 | 0.71857 | 0.006793 | 2142.5205 | 0.086733 | 139.6168 | 845.6816 | 207.5937 | 0.43664 | 0.054368 | 13.2332 | 26.9715 | 0.53988 | 0.0051038
846.3221 | 274.6154 | 0.54448 | 0.072303 | 16.9641 | 35.0914 | 0.72042 | 0.012895 | 2143.1184 | 0.086716 | 139.6785 | 846.3221 | 206.3992 | 0.40922 | 0.054342 | 12.7501 | 26.3744 | 0.54146 | 0.0096918
846.9625 | 273.141 | 0.5139 | 0.071626 | 16.3537 | 34.302 | 0.69792 | 0.008296 | 2144.3418 | 0.086631 | 139.7386 | 846.9625 | 205.4639 | 0.38657 | 0.053879 | 12.3017 | 25.8029 | 0.52499 | 0.0062405
847.603 | 271.7497 | 0.4797 | 0.071402 | 15.766 | 33.562 | 0.68054 | 0.005495 | 2144.6204 | 0.086526 | 139.797 | 847.603 | 204.7247 | 0.36139 | 0.053791 | 11.8774 | 25.2842 | 0.51269 | 0.0041397
848.2434 | 270.7228 | 0.45287 | 0.072245 | 15.2008 | 32.8585 | 0.65468 | 0.000873 | 2142.509 | 0.086577 | 139.8536 | 848.2434 | 204.1147 | 0.34145 | 0.05447 | 11.4608 | 24.7741 | 0.49361 | 0.00065821
848.8839 | 269.7162 | 0.42551 | 0.071594 | 14.6735 | 32.2005 | 0.6566 | 0.005086 | 2130.5042 | 0.086545 | 139.9086 | 848.8839 | 204.6583 | 0.32287 | 0.054325 | 11.1342 | 24.4335 | 0.49822 | 0.0038592
849.5242 | 268.8201 | 0.40315 | 0.071277 | 14.1582 | 31.5256 | 0.65343 | 0.009292 | 2108.2636 | 0.086462 | 139.9617 | 849.5242 | 206.4055 | 0.30955 | 0.054728 | 10.871 | 24.206 | 0.50171 | 0.0071346
850.1647 | 268.0703 | 0.38047 | 0.073631 | 13.6615 | 30.9072 | 0.64289 | 0.008036 | 2076.509 | 0.086547 | 140.013 | 850.1647 | 208.8487 | 0.29641 | 0.057365 | 10.6435 | 24.0792 | 0.50086 | 0.0062607
850.8051 | 267.1797 | 0.35787 | 0.071574 | 13.1915 | 30.2669 | 0.61991 | 0.010854 | 2037.1761 | 0.086536 | 140.0626 | 850.8051 | 212.275 | 0.28433 | 0.056866 | 10.4807 | 24.0471 | 0.49252 | 0.0086235
851.4456 | 266.0995 | 0.33847 | 0.072747 | 12.7248 | 29.6613 | 0.62584 | 0.010045 | 1994.2559 | 0.086353 | 140.1106 | 851.4456 | 216.4986 | 0.27538 | 0.059187 | 10.3529 | 24.1325 | 0.50919 | 0.0081726
852.086 | 265.1921 | 0.31888 | 0.073743 | 12.2513 | 29.0464 | 0.58929 | 0.002669 | 1954.9523 | 0.086459 | 140.1573 | 852.086 | 219.9015 | 0.26442 | 0.061149 | 10.1589 | 24.0857 | 0.48865 | 0.0022132
852.7264 | 264.2609 | 0.30045 | 0.073171 | 11.8125 | 28.4186 | 0.59018 | 0.009353 | 1921.1472 | 0.086468 | 140.2028 | 852.7264 | 223.0358 | 0.25358 | 0.061756 | 9.9697 | 23.9853 | 0.49811 | 0.0078939
853.3669 | 263.5793 | 0.2837 | 0.073703 | 11.3914 | 27.8459 | 0.58034 | 0.010835 | 1894.1345 | 0.086391 | 140.2473 | 853.3669 | 225.9046 | 0.24315 | 0.063168 | 9.7631 | 23.8658 | 0.49739 | 0.0092863
854.0073 | 262.9102 | 0.27009 | 0.074559 | 10.9841 | 27.3046 | 0.5677 | 0.005724 | 1872.1717 | 0.086292 | 140.2911 | 854.0073 | 228.3069 | 0.23455 | 0.064746 | 9.5384 | 23.7108 | 0.49298 | 0.0049706
854.6478 | 262.3859 | 0.25563 | 0.073555 | 10.6118 | 26.7821 | 0.57266 | 0.008289 | 1856.1421 | 0.086431 | 140.3342 | 854.6478 | 229.5223 | 0.22361 | 0.064342 | 9.2827 | 23.4276 | 0.50094 | 0.0072508
855.2881 | 261.9209 | 0.2443 | 0.075502 | 10.252 | 26.2822 | 0.56972 | 0.009215 | 1843.0746 | 0.086397 | 140.3771 | 855.2881 | 230.9005 | 0.21536 | 0.06656 | 9.0378 | 23.1695 | 0.50224 | 0.0081236
855.9286 | 261.5197 | 0.22998 | 0.075738 | 9.8913 | 25.7758 | 0.5447 | 0.008962 | 1831.0757 | 0.086298 | 140.4197 | 855.9286 | 232.3934 | 0.20437 | 0.067303 | 8.7896 | 22.9051 | 0.48404 | 0.0079639
856.569 | 260.8886 | 0.21964 | 0.076645 | 9.5463 | 25.2775 | 0.53783 | 0.007973 | 1821.0713 | 0.08627 | 140.4624 | 856.569 | 233.2538 | 0.19637 | 0.068526 | 8.5351 | 22.6 | 0.48086 | 0.0071285
857.2095 | 260.3687 | 0.20898 | 0.076771 | 9.1999 | 24.76 | 0.53662 | 0.002426 | 1812.9421 | 0.086337 | 140.5053 | 857.2095 | 233.722 | 0.18759 | 0.068914 | 8.2583 | 22.226 | 0.4817 | 0.0021777
857.8499 | 259.6209 | 0.19776 | 0.076378 | 8.8591 | 24.2776 | 0.52353 | 0.011147 | 1803.3125 | 0.086354 | 140.5487 | 857.8499 | 234.3222 | 0.17849 | 0.068935 | 7.9958 | 21.9118 | 0.47251 | 0.010061
858.4904 | 259.1035 | 0.18731 | 0.076257 | 8.5281 | 23.7961 | 0.5014 | 0.005555 | 1796.5504 | 0.086331 | 140.5927 | 858.4904 | 234.8709 | 0.16979 | 0.069125 | 7.7305 | 21.5706 | 0.4545 | 0.0050355
859.1308 | 258.4515 | 0.17761 | 0.077299 | 8.2182 | 23.297 | 0.51101 | 0.002671 | 1790.5789 | 0.086279 | 140.6375 | 859.1308 | 235.2792 | 0.16168 | 0.070369 | 7.4814 | 21.2083 | 0.4652 | 0.0024315
859.7712 | 257.7156 | 0.16899 | 0.078114 | 7.905 | 22.8406 | 0.50704 | 0.00137 | 1783.2693 | 0.086234 | 140.6833 | 859.7712 | 235.7686 | 0.1546 | 0.071462 | 7.2318 | 20.8955 | 0.46386 | 0.0012533
860.4117 | 257.0484 | 0.16111 | 0.078401 | 7.6087 | 22.3891 | 0.49512 | 0.011396 | 1778.1092 | 0.086292 | 140.7304 | 860.4117 | 235.7629 | 0.14777 | 0.071909 | 6.9786 | 20.5351 | 0.45412 | 0.010452
861.0521 | 256.4309 | 0.15427 | 0.078513 | 7.3171 | 21.9524 | 0.49224 | 0.003721 | 1774.3192 | 0.086317 | 140.7787 | 861.0521 | 235.7096 | 0.1418 | 0.072169 | 6.7258 | 20.1785 | 0.45247 | 0.0034203
861.6925 | 256.2272 | 0.14651 | 0.077996 | 7.0499 | 21.5697 | 0.46442 | 0.008069 | 1772.7247 | 0.086451 | 140.828 | 861.6925 | 235.4529 | 0.13464 | 0.071672 | 6.4784 | 19.8209 | 0.42676 | 0.0074148
862.3329 | 255.8024 | 0.14114 | 0.078173 | 6.7938 | 21.1668 | 0.46097 | 0.00516 | 1768.9125 | 0.086422 | 140.8783 | 862.3329 | 235.7328 | 0.13007 | 0.07204 | 6.2608 | 19.5061 | 0.4248 | 0.0047552
862.9734 | 255.4311 | 0.13619 | 0.079681 | 6.5271 | 20.7636 | 0.46711 | 0.010405 | 1766.9672 | 0.086471 | 140.9294 | 862.9734 | 235.5995 | 0.12562 | 0.073495 | 6.0204 | 19.1516 | 0.43084 | 0.0095972
863.6138 | 255.0857 | 0.12768 | 0.079413 | 6.2884 | 20.3922 | 0.45454 | 0.001574 | 1763.3543 | 0.086484 | 140.9811 | 863.6138 | 235.8141 | 0.11803 | 0.073413 | 5.8133 | 18.8516 | 0.4202 | 0.0014551
864.2543 | 254.6308 | 0.1228 | 0.07889 | 6.0454 | 19.9849 | 0.434 | 0.009274 | 1760.4217 | 0.086454 | 141.0333 | 864.2543 | 235.9551 | 0.11379 | 0.073104 | 5.602 | 18.5191 | 0.40217 | 0.0085938
864.8947 | 254.1051 | 0.12012 | 0.079226 | 5.8139 | 19.5886 | 0.43723 | 0.010831 | 1755.6484 | 0.08645 | 141.0858 | 864.8947 | 236.2081 | 0.11166 | 0.073646 | 5.4044 | 18.2089 | 0.40644 | 0.010068
865.5352 | 253.3161 | 0.1134 | 0.078705 | 5.5794 | 19.1903 | 0.45292 | 0.007615 | 1750.2449 | 0.086475 | 141.1385 | 865.5352 | 236.2203 | 0.10574 | 0.073393 | 5.2029 | 17.8952 | 0.42235 | 0.0071011
866.1755 | 252.5555 | 0.10895 | 0.077753 | 5.349 | 18.8234 | 0.43173 | 0.006678 | 1744.3836 | 0.086438 | 141.1913 | 866.1755 | 236.4931 | 0.10203 | 0.072808 | 5.0088 | 17.6262 | 0.40427 | 0.0062533
866.816 | 252.0133 | 0.10593 | 0.077962 | 5.1415 | 18.4353 | 0.42579 | 0.003021 | 1739.3559 | 0.086569 | 141.2441 | 866.816 | 236.3978 | 0.099363 | 0.073131 | 4.8229 | 17.293 | 0.39941 | 0.0028338
867.4564 | 251.4706 | 0.102 | 0.078383 | 4.9376 | 18.078 | 0.42041 | 0.006375 | 1736.2537 | 0.086514 | 141.2966 | 867.4564 | 236.5472 | 0.09595 | 0.073731 | 4.6446 | 17.0052 | 0.39546 | 0.0059967
868.0969 | 251.0999 | 0.097792 | 0.077932 | 4.7349 | 17.7364 | 0.39742 | -0.000734 | 1733.501 | 0.086507 | 141.3487 | 868.0969 | 236.68 | 0.092176 | 0.073457 | 4.463 | 16.7179 | 0.3746 | -0.00069185
868.7373 | 251.0452 | 0.095689 | 0.078278 | 4.5604 | 17.4185 | 0.40614 | 0.00666 | 1732.6884 | 0.08658 | 141.4003 | 868.7373 | 236.6271 | 0.090193 | 0.073782 | 4.2985 | 16.4182 | 0.38282 | 0.0062775
869.3777 | 251.0886 | 0.091769 | 0.078252 | 4.3879 | 17.1106 | 0.4086 | 0.008387 | 1731.7185 | 0.086617 | 141.4513 | 869.3777 | 236.7857 | 0.086542 | 0.073794 | 4.1379 | 16.1359 | 0.38533 | 0.0079092
870.0182 | 251.3598 | 0.091899 | 0.076717 | 4.2185 | 16.8207 | 0.38323 | 0.004472 | 1734.4972 | 0.086672 | 141.5014 | 870.0182 | 236.5937 | 0.0865 | 0.07221 | 3.9707 | 15.8326 | 0.36072 | 0.0042093
870.6586 | 251.7877 | 0.08885 | 0.077535 | 4.0665 | 16.5298 | 0.3992 | -0.000323 | 1738.225 | 0.08665 | 141.5507 | 870.6586 | 236.6298 | 0.083501 | 0.072867 | 3.8217 | 15.5347 | 0.37516 | -0.00030356
871.2991 | 252.6172 | 0.085985 | 0.077921 | 3.9216 | 16.2915 | 0.39528 | 0.002161 | 1746.1561 | 0.086723 | 141.599 | 871.2991 | 236.2147 | 0.080402 | 0.072862 | 3.667 | 15.2337 | 0.36961 | 0.0020207
871.9395 | 253.8647 | 0.082233 | 0.077422 | 3.7897 | 16.0465 | 0.37999 | 0.002841 | 1757.5971 | 0.086733 | 141.6466 | 871.9395 | 235.8885 | 0.07641 | 0.07194 | 3.5214 | 14.9103 | 0.35308 | 0.0026398
872.5799 | 255.5821 | 0.079632 | 0.078152 | 3.6656 | 15.8671 | 0.39242 | 0.004664 | 1773.4648 | 0.08674 | 141.6933 | 872.5799 | 235.4181 | 0.073349 | 0.071986 | 3.3764 | 14.6153 | 0.36146 | 0.004296
873.2203 | 257.8351 | 0.079612 | 0.076853 | 3.5552 | 15.7212 | 0.40195 | 0.00831 | 1793.1582 | 0.086858 | 141.7394 | 873.2203 | 234.6416 | 0.072451 | 0.06994 | 3.2354 | 14.307 | 0.36579 | 0.0075625
873.8608 | 260.4397 | 0.079215 | 0.077164 | 3.4523 | 15.5926 | 0.37288 | 0.003557 | 1816.6806 | 0.086902 | 141.7849 | 873.8608 | 233.8985 | 0.071142 | 0.0693 | 3.1005 | 14.0035 | 0.33488 | 0.0031945
874.5012 | 263.4753 | 0.077626 | 0.078326 | 3.3596 | 15.4993 | 0.37663 | 0.009215 | 1843.9522 | 0.086905 | 141.8297 | 874.5012 | 233.1904 | 0.068703 | 0.069323 | 2.9734 | 13.7178 | 0.33334 | 0.0081558
875.1417 | 267.0045 | 0.074742 | 0.078322 | 3.2647 | 15.3971 | 0.36842 | 0.009161 | 1874.1895 | 0.086941 | 141.8741 | 875.1417 | 232.4787 | 0.065077 | 0.068194 | 2.8426 | 13.4062 | 0.32078 | 0.0079764
875.7821 | 270.5561 | 0.074647 | 0.079466 | 3.1801 | 15.3434 | 0.38792 | 0.008155 | 1905.0793 | 0.087042 | 141.918 | 875.7821 | 231.554 | 0.063886 | 0.068011 | 2.7217 | 13.1315 | 0.332 | 0.0069794
876.4225 | 274.0994 | 0.071524 | 0.079614 | 3.0979 | 15.2479 | 0.38132 | 0.010069 | 1935.2648 | 0.087035 | 141.9615 | 876.4225 | 231.0173 | 0.060282 | 0.067101 | 2.611 | 12.8513 | 0.32139 | 0.0084864
877.0629 | 277.5454 | 0.071038 | 0.07898 | 3.0109 | 15.177 | 0.37185 | 0.009493 | 1963.6914 | 0.087027 | 142.0047 | 877.0629 | 230.6253 | 0.059029 | 0.065628 | 2.5019 | 12.6113 | 0.30898 | 0.0078882
877.7034 | 280.5282 | 0.072479 | 0.079956 | 2.9242 | 15.0711 | 0.3731 | 0.002123 | 1989.3409 | 0.087099 | 142.0475 | 877.7034 | 229.9799 | 0.059419 | 0.065549 | 2.3973 | 12.3555 | 0.30587 | 0.0017405
878.3438 | 283.228 | 0.069672 | 0.079106 | 2.8411 | 14.9473 | 0.37088 | 0.01041 | 2011.8974 | 0.087098 | 142.0902 | 878.3438 | 229.6616 | 0.056495 | 0.064145 | 2.3037 | 12.1203 | 0.30074 | 0.0084412
878.9843 | 285.5881 | 0.07083 | 0.078565 | 2.741 | 14.7961 | 0.37599 | 0.011914 | 2031.1447 | 0.087119 | 142.1327 | 878.9843 | 229.3943 | 0.056893 | 0.063106 | 2.2016 | 11.8847 | 0.30201 | 0.0095697
879.6247 | 287.2426 | 0.068659 | 0.078003 | 2.6516 | 14.5983 | 0.37429 | 0.007757 | 2046.1191 | 0.087123 | 142.1752 | 879.6247 | 229.0907 | 0.054759 | 0.062211 | 2.1148 | 11.6429 | 0.29851 | 0.0061866
880.2651 | 288.6108 | 0.068364 | 0.077807 | 2.5615 | 14.4236 | 0.36516 | 0.005636 | 2056.4398 | 0.087156 | 142.2175 | 880.2651 | 229.0101 | 0.054246 | 0.061739 | 2.0325 | 11.445 | 0.28975 | 0.0044721
880.9056 | 289.2115 | 0.065503 | 0.077219 | 2.461 | 14.1783 | 0.35827 | 0.005321 | 2062.8054 | 0.087218 | 142.2599 | 880.9056 | 228.6839 | 0.051794 | 0.061058 | 1.946 | 11.211 | 0.28329 | 0.0042074
881.546 | 289.5163 | 0.064314 | 0.076755 | 2.3663 | 13.9352 | 0.36155 | 0.010886 | 2066.75 | 0.087186 | 142.3021 | 881.546 | 228.6391 | 0.050791 | 0.060616 | 1.8687 | 11.005 | 0.28553 | 0.008597
882.1842 | 289.4959 | 0.062558 | 0.075235 | 2.2759 | 13.7021 | 0.3546 | 0.014853 | 2066.1463 | 0.087234 | 142.3439 | 882.1842 | 228.6318 | 0.049406 | 0.059417 | 1.7974 | 10.8213 | 0.28005 | 0.01173
882.8224 | 288.9305 | 0.062935 | 0.073736 | 2.1852 | 13.4124 | 0.33798 | 0.009967 | 2063.035 | 0.087158 | 142.3854 | 882.8224 | 228.7951 | 0.049836 | 0.058389 | 1.7304 | 10.6208 | 0.26763 | 0.0078926
883.4606 | 288.1373 | 0.061769 | 0.073467 | 2.0854 | 13.1436 | 0.3593 | 0.005302 | 2058.1939 | 0.087266 | 142.4265 | 883.4606 | 228.4847 | 0.048981 | 0.058257 | 1.6536 | 10.4225 | 0.28492 | 0.0042043
884.0988 | 287.3397 | 0.059066 | 0.071939 | 1.9974 | 12.8705 | 0.34036 | 0.010375 | 2051.0806 | 0.087329 | 142.4671 | 884.0988 | 228.5436 | 0.04698 | 0.057219 | 1.5887 | 10.2369 | 0.27072 | 0.008252
884.7369 | 286.4308 | 0.058653 | 0.070434 | 1.9113 | 12.5782 | 0.33784 | 0.005019 | 2046.4669 | 0.087261 | 142.507 | 884.7369 | 228.5771 | 0.046806 | 0.056208 | 1.5252 | 10.0377 | 0.2696 | 0.0040053
885.3751 | 285.5081 | 0.058233 | 0.069895 | 1.8249 | 12.3219 | 0.30356 | -0.000548 | 2039.7608 | 0.087266 | 142.5462 | 885.3751 | 228.6394 | 0.046634 | 0.055973 | 1.4614 | 9.8675 | 0.2431 | -0.00043885
886.0133 | 284.5782 | 0.055868 | 0.068091 | 1.7485 | 12.066 | 0.30612 | 0.009763 | 2032.4602 | 0.087245 | 142.5846 | 886.0133 | 228.8295 | 0.044924 | 0.054752 | 1.406 | 9.7023 | 0.24616 | 0.0078504
886.6515 | 283.7126 | 0.05283 | 0.066418 | 1.6721 | 11.8096 | 0.33533 | 0.006242 | 2026.4512 | 0.087337 | 142.6221 | 886.6515 | 228.6276 | 0.042573 | 0.053522 | 1.3475 | 9.5167 | 0.27022 | 0.0050301
887.2897 | 282.8763 | 0.054378 | 0.066361 | 1.5983 | 11.5499 | 0.31972 | 0.004748 | 2020.7967 | 0.087356 | 142.6586 | 887.2897 | 228.6008 | 0.043944 | 0.053628 | 1.2917 | 9.3338 | 0.25838 | 0.003837
887.9279 | 281.9529 | 0.051535 | 0.064903 | 1.5269 | 11.289 | 0.31715 | 0.009347 | 2013.7637 | 0.087296 | 142.6939 | 887.9279 | 228.8643 | 0.041832 | 0.052682 | 1.2394 | 9.1634 | 0.25744 | 0.0075871
888.566 | 281.1623 | 0.049046 | 0.064208 | 1.458 | 11.072 | 0.30607 | 0.002807 | 2007.7906 | 0.087326 | 142.7281 | 888.566 | 228.8784 | 0.039926 | 0.052268 | 1.1869 | 9.0131 | 0.24915 | 0.002285
889.2042 | 280.2875 | 0.0514 | 0.063368 | 1.3985 | 10.8197 | 0.30235 | 0.007162 | 2002.6035 | 0.087263 | 142.761 | 889.2042 | 228.9758 | 0.04199 | 0.051767 | 1.1425 | 8.839 | 0.247 | 0.0058509
889.8424 | 279.4324 | 0.051504 | 0.062575 | 1.339 | 10.6053 | 0.28956 | 0.002925 | 1995.9307 | 0.087308 | 142.7925 | 889.8424 | 228.9718 | 0.042203 | 0.051275 | 1.0972 | 8.6901 | 0.23727 | 0.0023968
890.4807 | 278.5574 | 0.048133 | 0.060742 | 1.2787 | 10.3771 | 0.28754 | -0.000378 | 1989.3725 | 0.087281 | 142.8225 | 890.4807 | 229.1262 | 0.039592 | 0.049963 | 1.0518 | 8.5357 | 0.23651 | -0.00031092
891.1188 | 277.4161 | 0.049094 | 0.060049 | 1.2192 | 10.1683 | 0.2753 | 0.000416 | 1981.6895 | 0.087435 | 142.8511 | 891.1188 | 228.7142 | 0.040475 | 0.049507 | 1.0051 | 8.3832 | 0.22697 | 0.00034297
891.757 | 276.4355 | 0.048009 | 0.058985 | 1.1694 | 9.9408 | 0.26304 | -0.000158 | 1973.8525 | 0.087389 | 142.8785 | 891.757 | 228.975 | 0.039766 | 0.048858 | 0.96862 | 8.2341 | 0.21788 | -0.00013087
892.3952 | 275.4046 | 0.048044 | 0.058068 | 1.1116 | 9.7056 | 0.2576 | 0.00379 | 1966.3791 | 0.087392 | 142.9048 | 892.3952 | 229.0223 | 0.039953 | 0.048288 | 0.92441 | 8.071 | 0.21421 | 0.0031517
893.0334 | 274.2977 | 0.045783 | 0.056489 | 1.0615 | 9.4969 | 0.27482 | 0.004311 | 1958.7888 | 0.087404 | 142.9303 | 893.0334 | 228.9968 | 0.038222 | 0.04716 | 0.88616 | 7.9285 | 0.22943 | 0.003599
893.6716 | 273.2493 | 0.043165 | 0.055038 | 1.0185 | 9.2917 | 0.27674 | 0.004593 | 1950.9856 | 0.087368 | 142.9551 | 893.6716 | 229.1682 | 0.036202 | 0.046159 | 0.85416 | 7.7927 | 0.23209 | 0.003852
894.3098 | 271.8807 | 0.042894 | 0.054153 | 0.96947 | 9.0661 | 0.25597 | 0.006006 | 1941.986 | 0.0874 | 142.9794 | 894.3098 | 229.0313 | 0.036134 | 0.045618 | 0.81668 | 7.6372 | 0.21563 | 0.0050594
894.9479 | 270.823 | 0.044408 | 0.053762 | 0.92588 | 8.8791 | 0.26299 | 0.003798 | 1934.2703 | 0.087438 | 143.0033 | 894.9479 | 228.9879 | 0.037548 | 0.045457 | 0.78285 | 7.5075 | 0.22236 | 0.0032113
895.5861 | 269.671 | 0.042862 | 0.05308 | 0.88568 | 8.6849 | 0.24732 | -0.006787 | 1926.3431 | 0.087404 | 143.027 | 895.5861 | 229.0791 | 0.03641 | 0.04509 | 0.75236 | 7.3777 | 0.21009 | -0.0057654
896.2243 | 268.7403 | 0.043253 | 0.051907 | 0.83864 | 8.4974 | 0.25037 | 0.005767 | 1919.9028 | 0.087494 | 143.0507 | 896.2243 | 228.8578 | 0.036834 | 0.044204 | 0.71418 | 7.2364 | 0.21321 | 0.0049111
896.8625 | 267.6421 | 0.04099 | 0.051457 | 0.80733 | 8.3039 | 0.2466 | 0.009707 | 1911.5728 | 0.087443 | 143.0746 | 896.8625 | 229.0875 | 0.035085 | 0.044044 | 0.69103 | 7.1077 | 0.21107 | 0.0083087
897.5007 | 266.8425 | 0.037399 | 0.049823 | 0.76779 | 8.118 | 0.2315 | 0.008662 | 1905.6348 | 0.087433 | 143.0988 | 897.5007 | 229.1803 | 0.03212 | 0.042791 | 0.65943 | 6.9722 | 0.19883 | 0.0074394
898.1389 | 265.976 | 0.039115 | 0.048932 | 0.73174 | 7.9461 | 0.23114 | 0.013065 | 1899.5588 | 0.087462 | 143.1235 | 898.1389 | 229.1287 | 0.033696 | 0.042153 | 0.63037 | 6.8453 | 0.19912 | 0.011255
898.777 | 265.4978 | 0.036491 | 0.049429 | 0.70204 | 7.7886 | 0.22397 | 0.007032 | 1896.0907 | 0.087476 | 143.1489 | 898.777 | 229.1413 | 0.031494 | 0.04266 | 0.6059 | 6.722 | 0.1933 | 0.0060691
899.4152 | 264.9457 | 0.036673 | 0.048253 | 0.67341 | 7.6403 | 0.22928 | 0.007803 | 1892.2544 | 0.087515 | 143.1751 | 899.4152 | 229.0681 | 0.031707 | 0.041719 | 0.58222 | 6.6057 | 0.19823 | 0.0067464
900.0535 | 264.4374 | 0.036906 | 0.046101 | 0.64133 | 7.4895 | 0.23004 | 0.005167 | 1888.2204 | 0.087569 | 143.2023 | 900.0535 | 229.0177 | 0.031963 | 0.039926 | 0.55542 | 6.4863 | 0.19923 | 0.0044749
900.6916 | 264.1557 | 0.037749 | 0.045677 | 0.6136 | 7.3409 | 0.21834 | 0.005013 | 1886.3317 | 0.087569 | 143.2307 | 900.6916 | 229.0491 | 0.032732 | 0.039606 | 0.53205 | 6.3653 | 0.18932 | 0.0043468
901.3298 | 263.8718 | 0.037011 | 0.045631 | 0.58839 | 7.2061 | 0.22418 | 0.001899 | 1884.9412 | 0.087484 | 143.2603 | 901.3298 | 229.2401 | 0.032154 | 0.039642 | 0.51117 | 6.2604 | 0.19476 | 0.0016498
901.968 | 263.6555 | 0.035864 | 0.045746 | 0.56117 | 7.0644 | 0.21701 | 0.008976 | 1883.2352 | 0.08755 | 143.2912 | 901.968 | 229.1375 | 0.031169 | 0.039757 | 0.4877 | 6.1395 | 0.18859 | 0.0078009
902.6062 | 263.3658 | 0.034663 | 0.044517 | 0.53948 | 6.9215 | 0.20698 | 0.006524 | 1880.5772 | 0.087529 | 143.3233 | 902.6062 | 229.3154 | 0.030181 | 0.038761 | 0.46973 | 6.0266 | 0.18022 | 0.0056805
903.2444 | 263.0726 | 0.034836 | 0.043173 | 0.51728 | 6.7887 | 0.21511 | 0.009982 | 1878.9058 | 0.087639 | 143.3568 | 903.2444 | 229.0303 | 0.030328 | 0.037586 | 0.45034 | 5.9102 | 0.18727 | 0.0086903
903.8826 | 262.9653 | 0.034958 | 0.041612 | 0.49325 | 6.6648 | 0.18809 | 0.010502 | 1877.2178 | 0.087604 | 143.3916 | 903.8826 | 229.289 | 0.030481 | 0.036283 | 0.43008 | 5.8113 | 0.164 | 0.0091571
904.5208 | 262.8197 | 0.032782 | 0.042193 | 0.47516 | 6.5431 | 0.2047 | 0.004216 | 1876.2407 | 0.087577 | 143.4277 | 904.5208 | 229.4098 | 0.028615 | 0.036829 | 0.41476 | 5.7113 | 0.17868 | 0.0036801
905.1589 | 262.6791 | 0.033362 | 0.041613 | 0.45696 | 6.4127 | 0.19979 | 0.007455 | 1874.6467 | 0.087662 | 143.4653 | 905.1589 | 229.3186 | 0.029125 | 0.036328 | 0.39892 | 5.5983 | 0.17442 | 0.0065082
905.7971 | 262.4499 | 0.031645 | 0.040519 | 0.43366 | 6.2897 | 0.20111 | 0.002524 | 1872.8658 | 0.087631 | 143.5044 | 905.7971 | 229.4803 | 0.02767 | 0.035429 | 0.37919 | 5.4996 | 0.17584 | 0.0022069
906.4353 | 262.3358 | 0.032527 | 0.04121 | 0.41591 | 6.1722 | 0.19846 | 0.005069 | 1872.2945 | 0.087633 | 143.5449 | 906.4353 | 229.5098 | 0.028457 | 0.036053 | 0.36387 | 5.3998 | 0.17362 | 0.0044347
907.0735 | 262.1915 | 0.033818 | 0.038813 | 0.40284 | 6.0524 | 0.1924 | 0.008516 | 1871.1858 | 0.087633 | 143.5869 | 907.0735 | 229.588 | 0.029613 | 0.033987 | 0.35275 | 5.2998 | 0.16847 | 0.007457
907.7117 | 262.1603 | 0.03225 | 0.038835 | 0.38537 | 5.942 | 0.20164 | 0.001986 | 1870.939 | 0.087659 | 143.6305 | 907.7117 | 229.5912 | 0.028243 | 0.03401 | 0.3375 | 5.2038 | 0.17659 | 0.0017393
908.3499 | 261.9964 | 0.032466 | 0.038674 | 0.37032 | 5.8298 | 0.16779 | 0.00157 | 1869.4819 | 0.08769 | 143.6756 | 908.3499 | 229.6173 | 0.028454 | 0.033894 | 0.32455 | 5.1094 | 0.14706 | 0.001376
908.988 | 261.8448 | 0.032423 | 0.038183 | 0.35356 | 5.7159 | 0.19348 | 0.004323 | 1868.2549 | 0.087772 | 143.7224 | 908.988 | 229.4963 | 0.028417 | 0.033466 | 0.30988 | 5.0098 | 0.16958 | 0.0037889
909.6262 | 261.5674 | 0.031156 | 0.037605 | 0.34246 | 5.6022 | 0.17947 | -0.002216 | 1866.232 | 0.087776 | 143.7709 | 909.6262 | 229.5679 | 0.027344 | 0.033004 | 0.30056 | 4.9169 | 0.15752 | -0.0019449
910.2645 | 261.3323 | 0.029783 | 0.037128 | 0.3273 | 5.495 | 0.19665 | -0.002862 | 1865.4349 | 0.087745 | 143.821 | 910.2645 | 229.6221 | 0.026169 | 0.032623 | 0.28758 | 4.8282 | 0.17279 | -0.0025147
910.9026 | 261.0672 | 0.030094 | 0.036347 | 0.31188 | 5.3993 | 0.19002 | 0.000568 | 1863.8155 | 0.087717 | 143.8725 | 910.9026 | 229.7429 | 0.026483 | 0.031986 | 0.27446 | 4.7515 | 0.16722 | 0.00049985
911.5408 | 260.796 | 0.030556 | 0.036245 | 0.29665 | 5.2924 | 0.17036 | -0.001447 | 1862.1483 | 0.087857 | 143.9249 | 911.5408 | 229.4284 | 0.026881 | 0.031886 | 0.26097 | 4.6559 | 0.14987 | -0.001273
912.179 | 260.3889 | 0.030351 | 0.036185 | 0.28574 | 5.1884 | 0.19377 | 0.000827 | 1860.3272 | 0.087768 | 143.9774 | 912.179 | 229.6107 | 0.026763 | 0.031908 | 0.25197 | 4.5751 | 0.17086 | 0.00072925
912.809 | 260.1404 | 0.030657 | 0.03462 | 0.27295 | 5.0944 | 0.18428 | -0.000764 | 1857.6124 | 0.087781 | 144.0289 | 912.809 | 229.7754 | 0.027079 | 0.030579 | 0.24109 | 4.4997 | 0.16277 | -0.00067482
913.439 | 259.7522 | 0.026573 | 0.035066 | 0.26321 | 4.9981 | 0.18218 | -0.002337 | 1855.6925 | 0.087807 | 144.0794 | 913.439 | 229.6807 | 0.023497 | 0.031006 | 0.23274 | 4.4195 | 0.16109 | -0.0020664
914.069 | 259.3783 | 0.03014 | 0.034606 | 0.25129 | 4.8865 | 0.171 | -0.008889 | 1853.1959 | 0.087835 | 144.1283 | 914.069 | 229.6637 | 0.026687 | 0.030642 | 0.2225 | 4.3267 | 0.15141 | -0.0078707
914.699 | 259.0299 | 0.028093 | 0.034189 | 0.24022 | 4.7942 | 0.16605 | -0.00114 | 1849.2921 | 0.087829 | 144.1751 | 914.699 | 229.9306 | 0.024937 | 0.030348 | 0.21324 | 4.2557 | 0.1474 | -0.0010119
915.329 | 258.6727 | 0.028444 | 0.03337 | 0.23067 | 4.6929 | 0.1594 | 0.003456 | 1846.2415 | 0.087922 | 144.219 | 915.329 | 229.8208 | 0.025271 | 0.029648 | 0.20494 | 4.1695 | 0.14162 | 0.0030705
915.9589 | 258.3232 | 0.025887 | 0.03308 | 0.22218 | 4.6069 | 0.1777 | -0.001246 | 1844.57 | 0.087905 | 144.2596 | 915.9589 | 229.8267 | 0.023031 | 0.029431 | 0.19767 | 4.0987 | 0.1581 | -0.0011085
916.5889 | 257.7236 | 0.02746 | 0.032447 | 0.21538 | 4.5184 | 0.15915 | 0.005838 | 1840.999 | 0.087872 | 144.2962 | 916.5889 | 229.883 | 0.024494 | 0.028942 | 0.19211 | 4.0303 | 0.14196 | 0.0052074
917.2189 | 257.3288 | 0.024288 | 0.031441 | 0.20556 | 4.4248 | 0.16326 | 0.001593 | 1838.5855 | 0.087863 | 144.3283 | 917.2189 | 229.9063 | 0.0217 | 0.02809 | 0.18366 | 3.9533 | 0.14587 | 0.0014232
917.8489 | 256.911 | 0.026624 | 0.031173 | 0.19387 | 4.3418 | 0.16624 | 0.000524 | 1835.4576 | 0.087937 | 144.3552 | 917.8489 | 229.7729 | 0.023812 | 0.02788 | 0.17339 | 3.8832 | 0.14868 | 0.00046865
918.4789 | 256.5124 | 0.026051 | 0.031465 | 0.18817 | 4.2494 | 0.17432 | 0.002608 | 1832.6903 | 0.087955 | 144.3763 | 918.4789 | 229.7492 | 0.023333 | 0.028182 | 0.16854 | 3.8061 | 0.15613 | 0.0023359
919.1089 | 256.0843 | 0.0241 | 0.030505 | 0.1836 | 4.1696 | 0.16086 | -0.000308 | 1829.1998 | 0.087927 | 144.3911 | 919.1089 | 229.9016 | 0.021636 | 0.027386 | 0.16483 | 3.7433 | 0.14441 | -0.00027651
919.7389 | 255.621 | 0.026281 | 0.029712 | 0.17681 | 4.0915 | 0.13663 | 0.000379 | 1825.8303 | 0.087906 | 144.3989 | 919.7389 | 229.9758 | 0.023644 | 0.026731 | 0.15907 | 3.681 | 0.12292 | 0.00034098
920.3689 | 255.1415 | 0.024278 | 0.029908 | 0.16778 | 4.0017 | 0.16001 | 0.00945 | 1822.0685 | 0.087905 | 144.3993 | 920.3689 | 230.0224 | 0.021888 | 0.026964 | 0.15126 | 3.6077 | 0.14426 | 0.0085196
920.9988 | 254.6931 | 0.024058 | 0.028449 | 0.16428 | 3.9187 | 0.15222 | 0.001983 | 1816.8759 | 0.088042 | 144.3925 | 920.9988 | 229.9044 | 0.021716 | 0.02568 | 0.14829 | 3.5373 | 0.13741 | 0.00179
921.6288 | 254.2018 | 0.023582 | 0.029312 | 0.15491 | 3.8535 | 0.12968 | -0.001164 | 1813.4888 | 0.087994 | 144.3794 | 921.6288 | 229.9948 | 0.021336 | 0.026521 | 0.14016 | 3.4866 | 0.11733 | -0.0010532
922.2589 | 253.9883 | 0.024483 | 0.027014 | 0.1508 | 3.7873 | 0.14739 | 0.005449 | 1811.504 | 0.088059 | 144.3612 | 922.2589 | 229.8544 | 0.022157 | 0.024447 | 0.13647 | 3.4274 | 0.13339 | 0.0049312
922.8889 | 253.6323 | 0.023955 | 0.028564 | 0.14122 | 3.703 | 0.14186 | 0.001403 | 1809.2496 | 0.088023 | 144.3387 | 922.8889 | 229.8765 | 0.021711 | 0.025889 | 0.12799 | 3.3562 | 0.12857 | 0.0012716
923.5189 | 253.3919 | 0.022201 | 0.028338 | 0.139 | 3.6467 | 0.13916 | 0.004394 | 1807.0178 | 0.088068 | 144.3131 | 923.5189 | 229.7819 | 0.020132 | 0.025698 | 0.12605 | 3.3069 | 0.12619 | 0.0039846
924.1488 | 253.0305 | 0.022077 | 0.026956 | 0.13112 | 3.5643 | 0.13023 | 0.006476 | 1804.9232 | 0.088088 | 144.2852 | 924.1488 | 229.6244 | 0.020035 | 0.024462 | 0.11899 | 3.2346 | 0.11818 | 0.0058769
924.7788 | 252.6145 | 0.021545 | 0.027912 | 0.12836 | 3.4947 | 0.12583 | 0.010597 | 1802.5175 | 0.08809 | 144.2561 | 924.7788 | 229.5026 | 0.019574 | 0.025358 | 0.11662 | 3.175 | 0.11432 | 0.0096275
925.4088 | 252.1418 | 0.020451 | 0.026965 | 0.12238 | 3.424 | 0.13162 | 0.001565 | 1799.698 | 0.088025 | 144.2268 | 925.4088 | 229.5539 | 0.018619 | 0.024549 | 0.11141 | 3.1172 | 0.11983 | 0.0014248
926.0388 | 251.7922 | 0.021039 | 0.026338 | 0.11592 | 3.3639 | 0.12381 | 0.00364 | 1798.1666 | 0.088027 | 144.1983 | 926.0388 | 229.3815 | 0.019166 | 0.023994 | 0.1056 | 3.0645 | 0.11279 | 0.003316
926.6688 | 251.4077 | 0.020658 | 0.025911 | 0.11414 | 3.298 | 0.14378 | -0.000636 | 1795.353 | 0.088069 | 144.1716 | 926.6688 | 229.2382 | 0.018836 | 0.023626 | 0.10407 | 3.0071 | 0.13111 | -0.00057992
927.2988 | 251.0523 | 0.022981 | 0.025577 | 0.11268 | 3.2354 | 0.13315 | 0.00417 | 1792.5942 | 0.088151 | 144.1476 | 927.2988 | 229.0133 | 0.020964 | 0.023332 | 0.10279 | 2.9514 | 0.12146 | 0.0038039
927.9288 | 250.6591 | 0.020812 | 0.025625 | 0.10259 | 3.166 | 0.13384 | 0.005356 | 1789.3276 | 0.088092 | 144.1274 | 927.9288 | 229.1944 | 0.01903 | 0.023431 | 0.093807 | 2.8949 | 0.12238 | 0.0048973
928.5588 | 250.2989 | 0.020836 | 0.02551 | 0.09882 | 3.1102 | 0.11639 | 0.006315 | 1786.8089 | 0.08823 | 144.112 | 928.5588 | 228.8034 | 0.019047 | 0.023319 | 0.090333 | 2.8431 | 0.10639 | 0.0057727
929.1887 | 249.6432 | 0.020757 | 0.024303 | 0.098138 | 3.0421 | 0.12049 | 0.012023 | 1782.2781 | 0.088211 | 144.1023 | 929.1887 | 228.82 | 0.019026 | 0.022276 | 0.089952 | 2.7883 | 0.11044 | 0.01102
929.8187 | 249.2789 | 0.021443 | 0.025393 | 0.093303 | 2.9798 | 0.12006 | 0.003656 | 1780.3262 | 0.088282 | 144.0994 | 929.8187 | 228.5461 | 0.01966 | 0.023281 | 0.085543 | 2.732 | 0.11007 | 0.0033519
930.4487 | 248.6889 | 0.019423 | 0.024024 | 0.092283 | 2.9344 | 0.12528 | 0.006167 | 1776.687 | 0.088228 | 144.1042 | 930.4487 | 228.6204 | 0.017856 | 0.022085 | 0.084836 | 2.6976 | 0.11517 | 0.0056693
931.0787 | 248.1894 | 0.016843 | 0.023732 | 0.087046 | 2.8628 | 0.12653 | 0.002331 | 1773.4623 | 0.088283 | 144.1163 | 931.0787 | 228.454 | 0.015504 | 0.021845 | 0.080124 | 2.6351 | 0.11647 | 0.0021456
931.7087 | 247.7012 | 0.018495 | 0.02431 | 0.085755 | 2.8034 | 0.12155 | 0.000548 | 1769.8237 | 0.088322 | 144.1347 | 931.7087 | 228.4004 | 0.017054 | 0.022416 | 0.079073 | 2.585 | 0.11208 | 0.0005053
932.3387 | 247.0484 | 0.019738 | 0.023474 | 0.079717 | 2.7512 | 0.11033 | 0.00171 | 1765.5365 | 0.08837 | 144.1585 | 932.3387 | 228.2666 | 0.018237 | 0.021689 | 0.073657 | 2.542 | 0.10194 | 0.00158
932.9687 | 246.3386 | 0.018021 | 0.02291 | 0.074664 | 2.6965 | 0.11611 | 0.001506 | 1759.7001 | 0.08835 | 144.1867 | 932.9687 | 228.46 | 0.016713 | 0.021247 | 0.069245 | 2.5008 | 0.10769 | 0.0013967
933.5987 | 245.2976 | 0.019109 | 0.021814 | 0.07415 | 2.6446 | 0.12192 | -0.003017 | 1753.0598 | 0.088255 | 144.2182 | 933.5987 | 228.6541 | 0.017812 | 0.020334 | 0.069119 | 2.4652 | 0.11364 | -0.0028123
934.2287 | 244.4427 | 0.019279 | 0.022804 | 0.070996 | 2.5868 | 0.11864 | 0.002439 | 1745.7932 | 0.088432 | 144.252 | 934.2287 | 228.4006 | 0.018014 | 0.021307 | 0.066337 | 2.417 | 0.11085 | 0.0022789
934.8586 | 243.7598 | 0.017965 | 0.021729 | 0.071987 | 2.5347 | 0.11195 | 0.00745 | 1740.0945 | 0.088423 | 144.2873 | 934.8586 | 228.5886 | 0.016847 | 0.020377 | 0.067507 | 2.377 | 0.10498 | 0.0069863
935.4886 | 243.1492 | 0.016652 | 0.022445 | 0.067312 | 2.4731 | 0.10186 | 0.00454 | 1735.7833 | 0.088328 | 144.3229 | 935.4886 | 228.8833 | 0.015675 | 0.021128 | 0.063363 | 2.328 | 0.095888 | 0.0042736
936.1187 | 242.6238 | 0.016458 | 0.022036 | 0.065466 | 2.4233 | 0.12183 | 0.000944 | 1732.635 | 0.088437 | 144.3579 | 936.1187 | 228.5776 | 0.015505 | 0.02076 | 0.061676 | 2.283 | 0.11477 | 0.00088935
936.7487 | 242.2045 | 0.017438 | 0.022272 | 0.061571 | 2.3804 | 0.12194 | -0.001751 | 1728.8653 | 0.088352 | 144.3912 | 936.7487 | 228.9526 | 0.016484 | 0.021053 | 0.058202 | 2.2502 | 0.11527 | -0.0016552
937.3786 | 241.6254 | 0.016172 | 0.020797 | 0.059387 | 2.3294 | 0.10674 | -0.000152 | 1725.0106 | 0.088453 | 144.422 | 937.3786 | 228.7021 | 0.015307 | 0.019685 | 0.056211 | 2.2048 | 0.10103 | -0.00014387
938.0086 | 241.1715 | 0.017455 | 0.021153 | 0.05766 | 2.2795 | 0.10174 | 0.004385 | 1722.0983 | 0.088485 | 144.4492 | 938.0086 | 228.6189 | 0.016546 | 0.020052 | 0.054659 | 2.1608 | 0.096444 | 0.0041568
938.6386 | 240.6148 | 0.01619 | 0.020837 | 0.056455 | 2.24 | 0.10863 | 0.002818 | 1718.9682 | 0.088438 | 144.4717 | 938.6386 | 228.6649 | 0.015386 | 0.019802 | 0.053651 | 2.1288 | 0.10324 | 0.002678
939.2686 | 240.2409 | 0.016301 | 0.022312 | 0.053206 | 2.2031 | 0.095702 | -0.001471 | 1715.9752 | 0.088417 | 144.4887 | 939.2686 | 228.7883 | 0.015524 | 0.021248 | 0.05067 | 2.0981 | 0.09114 | -0.0014009
939.8986 | 240.0692 | 0.015798 | 0.01992 | 0.051472 | 2.1541 | 0.10093 | 0.000133 | 1714.6338 | 0.088447 | 144.499 | 939.8986 | 228.7428 | 0.015053 | 0.01898 | 0.049044 | 2.0525 | 0.096167 | 0.00012673
940.5286 | 239.4702 | 0.016469 | 0.021349 | 0.051051 | 2.106 | 0.12201 | -0.003025 | 1710.7362 | 0.088543 | 144.502 | 940.5286 | 228.4495 | 0.015711 | 0.020366 | 0.048702 | 2.0091 | 0.1164 | -0.0028858
941.1586 | 238.8865 | 0.017239 | 0.020291 | 0.046937 | 2.0703 | 0.11076 | 0.000856 | 1706.3572 | 0.088525 | 144.4981 | 941.1586 | 228.5165 | 0.016491 | 0.01941 | 0.044899 | 1.9805 | 0.10595 | 0.00081884
941.7886 | 238.183 | 0.014642 | 0.020287 | 0.049667 | 2.0257 | 0.097258 | -2.9e-05 | 1702.2027 | 0.088514 | 144.4883 | 941.7886 | 228.4132 | 0.014041 | 0.019455 | 0.04763 | 1.9426 | 0.093269 | -2.781e-05
942.4186 | 237.689 | 0.014748 | 0.020849 | 0.044164 | 1.9828 | 0.1053 | 0.002945 | 1698.1307 | 0.088555 | 144.4737 | 942.4186 | 228.357 | 0.014169 | 0.02003 | 0.04243 | 1.905 | 0.10116 | 0.0028294
943.0485 | 237.0571 | 0.013476 | 0.019834 | 0.045637 | 1.9464 | 0.096892 | -0.004056 | 1694.0871 | 0.088504 | 144.4554 | 943.0485 | 228.3968 | 0.012984 | 0.019109 | 0.04397 | 1.8753 | 0.093352 | -0.0039078
943.6785 | 236.5975 | 0.015113 | 0.01992 | 0.041181 | 1.9096 | 0.092459 | 0.002997 | 1690.147 | 0.088618 | 144.4343 | 943.6785 | 228.1572 | 0.014574 | 0.019209 | 0.039712 | 1.8415 | 0.089161 | 0.0028901
944.3085 | 235.8407 | 0.014298 | 0.018947 | 0.04344 | 1.8616 | 0.080093 | -0.003327 | 1685.5325 | 0.088579 | 144.4114 | 944.3085 | 228.1147 | 0.01383 | 0.018326 | 0.042017 | 1.8006 | 0.077469 | -0.003218
944.9385 | 235.077 | 0.012872 | 0.019235 | 0.043526 | 1.8289 | 0.07226 | -0.004663 | 1680.9581 | 0.088675 | 144.3879 | 944.9385 | 227.7094 | 0.012469 | 0.018632 | 0.042162 | 1.7716 | 0.069995 | -0.0045169
945.5685 | 234.5148 | 0.015321 | 0.018416 | 0.038802 | 1.7938 | 0.070092 | 0.008052 | 1676.844 | 0.088636 | 144.3648 | 945.5685 | 227.7867 | 0.014881 | 0.017888 | 0.037689 | 1.7424 | 0.068081 | 0.007821
946.1985 | 233.8741 | 0.016241 | 0.018181 | 0.037431 | 1.7542 | 0.087755 | -0.002594 | 1672.3829 | 0.088624 | 144.3431 | 946.1985 | 227.7681 | 0.015817 | 0.017706 | 0.036454 | 1.7084 | 0.085464 | -0.0025263
946.8285 | 233.4061 | 0.01364 | 0.019046 | 0.04001 | 1.7159 | 0.10712 | -0.001442 | 1669.4232 | 0.088768 | 144.3237 | 946.8285 | 227.3156 | 0.013284 | 0.018549 | 0.038966 | 1.6711 | 0.10433 | -0.0014044
947.4585 | 232.8976 | 0.014916 | 0.018331 | 0.035807 | 1.6771 | 0.085105 | -0.003926 | 1665.3184 | 0.088578 | 144.3078 | 947.4585 | 227.8416 | 0.014592 | 0.017933 | 0.03503 | 1.6407 | 0.083257 | -0.0038408
948.0884 | 232.4372 | 0.012669 | 0.017759 | 0.034365 | 1.6588 | 0.057414 | 0.002442 | 1662.1666 | 0.08872 | 144.2965 | 948.0884 | 227.4389 | 0.012397 | 0.017377 | 0.033626 | 1.6232 | 0.056179 | 0.0023895
948.7184 | 231.6424 | 0.013994 | 0.017269 | 0.034701 | 1.6183 | 0.087548 | 0.002924 | 1656.1975 | 0.088707 | 144.2906 | 948.7184 | 227.5012 | 0.013744 | 0.01696 | 0.034081 | 1.5894 | 0.085983 | 0.0028717
949.3484 | 230.8357 | 0.013077 | 0.017965 | 0.032334 | 1.5817 | 0.081401 | 0.001369 | 1651.2966 | 0.088711 | 144.2913 | 949.3484 | 227.3732 | 0.012881 | 0.017696 | 0.031849 | 1.5579 | 0.08018 | 0.0013485
949.9785 | 229.9937 | 0.012973 | 0.017095 | 0.031547 | 1.5432 | 0.070722 | -0.004723 | 1646.0663 | 0.088773 | 144.2996 | 949.9785 | 227.1174 | 0.012811 | 0.016881 | 0.031152 | 1.5239 | 0.069838 | -0.0046639
950.6085 | 229.3599 | 0.013431 | 0.01739 | 0.029635 | 1.5215 | 0.089517 | 0.000277 | 1641.0396 | 0.088753 | 144.3162 | 950.6085 | 227.2638 | 0.013308 | 0.017231 | 0.029364 | 1.5076 | 0.088699 | 0.00027447
951.2384 | 228.7074 | 0.01098 | 0.018096 | 0.031544 | 1.4866 | 0.09187 | 0.000959 | 1636.8638 | 0.088773 | 144.3406 | 951.2384 | 227.1817 | 0.010907 | 0.017975 | 0.031334 | 1.4767 | 0.091257 | 0.0009526
951.8684 | 228.1661 | 0.013018 | 0.0174 | 0.032329 | 1.459 | 0.078929 | 0.002095 | 1633.5956 | 0.088827 | 144.372 | 951.8684 | 227.0085 | 0.012952 | 0.017312 | 0.032165 | 1.4516 | 0.078529 | 0.0020844
952.4984 | 227.584 | 0.011128 | 0.016897 | 0.030477 | 1.4232 | 0.072781 | 0.005283 | 1629.3939 | 0.088829 | 144.4093 | 952.4984 | 227.0673 | 0.011103 | 0.016859 | 0.030408 | 1.42 | 0.072616 | 0.005271
953.1284 | 227.0951 | 0.012448 | 0.017618 | 0.02586 | 1.3997 | 0.063809 | -0.001961 | 1626.3582 | 0.088834 | 144.4518 | 953.1284 | 227.0584 | 0.012446 | 0.017615 | 0.025856 | 1.3995 | 0.063799 | -0.0019607
953.7584 | 226.5745 | 0.01188 | 0.01672 | 0.025717 | 1.3688 | 0.065598 | 0.002608 | 1623.0471 | 0.088924 | 144.4985 | 953.7584 | 226.8429 | 0.011894 | 0.01674 | 0.025747 | 1.3705 | 0.065676 | 0.0026111
954.3884 | 225.9697 | 0.012589 | 0.016946 | 0.029388 | 1.3457 | 0.065113 | -0.000761 | 1619.887 | 0.088931 | 144.5486 | 954.3884 | 226.7383 | 0.012632 | 0.017004 | 0.029488 | 1.3503 | 0.065334 | -0.00076359
955.0184 | 225.4931 | 0.00998 | 0.016688 | 0.026937 | 1.3168 | 0.072562 | -0.001202 | 1617.115 | 0.089006 | 144.6012 | 955.0184 | 226.541 | 0.010026 | 0.016766 | 0.027062 | 1.3229 | 0.072899 | -0.0012076
955.6484 | 225.1594 | 0.011503 | 0.01674 | 0.025275 | 1.2907 | 0.073224 | -0.000742 | 1615.8455 | 0.088916 | 144.6553 | 955.6484 | 226.6966 | 0.011582 | 0.016854 | 0.025448 | 1.2995 | 0.073724 | -0.00074707
956.2783 | 224.9423 | 0.012822 | 0.01671 | 0.024928 | 1.2737 | 0.065936 | -0.000282 | 1614.3129 | 0.088917 | 144.7101 | 956.2783 | 226.7765 | 0.012927 | 0.016846 | 0.025131 | 1.284 | 0.066474 | -0.0002843
956.9083 | 224.6139 | 0.009856 | 0.016508 | 0.023528 | 1.2416 | 0.074528 | 0.002219 | 1613.0184 | 0.089026 | 144.7647 | 956.9083 | 226.4359 | 0.009936 | 0.016642 | 0.023719 | 1.2517 | 0.075133 | 0.002237
957.5383 | 224.3915 | 0.012501 | 0.016386 | 0.023113 | 1.2106 | 0.066113 | -0.007968 | 1612.1783 | 0.089022 | 144.8182 | 957.5383 | 226.422 | 0.012614 | 0.016534 | 0.023322 | 1.2215 | 0.066711 | -0.0080401
958.1683 | 224.0332 | 0.011319 | 0.016595 | 0.024408 | 1.2022 | 0.062783 | -1.1e-05 | 1611.6404 | 0.089044 | 144.8697 | 958.1683 | 226.1613 | 0.011427 | 0.016753 | 0.02464 | 1.2136 | 0.063379 | -1.1104e-05
958.7983 | 223.8414 | 0.011768 | 0.016383 | 0.023131 | 1.1723 | 0.073061 | -0.002933 | 1610.2269 | 0.089079 | 144.9184 | 958.7983 | 226.1538 | 0.01189 | 0.016552 | 0.02337 | 1.1844 | 0.073816 | -0.0029633
959.4283 | 223.7101 | 0.011145 | 0.016077 | 0.02073 | 1.1514 | 0.05784 | -0.001552 | 1609.2476 | 0.08911 | 144.9633 | 959.4283 | 226.1483 | 0.011266 | 0.016252 | 0.020956 | 1.1639 | 0.05847 | -0.0015689
960.0583 | 223.621 | 0.010012 | 0.015489 | 0.022083 | 1.1358 | 0.073797 | -0.006285 | 1607.2087 | 0.08911 | 145.0035 | 960.0583 | 226.4087 | 0.010137 | 0.015682 | 0.022358 | 1.15 | 0.074717 | -0.0063633
960.6883 | 223.4873 | 0.009744 | 0.016476 | 0.022581 | 1.1174 | 0.073306 | 0.003496 | 1605.3218 | 0.089196 | 145.0385 | 960.6883 | 226.3742 | 0.0098699 | 0.016689 | 0.022873 | 1.1319 | 0.074253 | 0.0035412
961.3183 | 222.9187 | 0.011618 | 0.015726 | 0.020377 | 1.0926 | 0.076222 | 0.001476 | 1600.3487 | 0.089147 | 145.0687 | 961.3183 | 226.6716 | 0.011814 | 0.015991 | 0.02072 | 1.111 | 0.077505 | 0.0015008
961.9482 | 222.0896 | 0.009413 | 0.015336 | 0.020607 | 1.0691 | 0.058731 | -0.003499 | 1594.2852 | 0.089133 | 145.0946 | 961.9482 | 226.764 | 0.0096111 | 0.015659 | 0.021041 | 1.0916 | 0.059967 | -0.0035726
962.5782 | 221.3004 | 0.009453 | 0.016661 | 0.019191 | 1.0462 | 0.064478 | 0.000819 | 1588.4112 | 0.089142 | 145.1168 | 962.5782 | 226.8065 | 0.0096882 | 0.017076 | 0.019668 | 1.0723 | 0.066082 | 0.00083938
963.2083 | 220.3125 | 0.008929 | 0.015329 | 0.019947 | 1.0214 | 0.067961 | 0.001962 | 1581.2272 | 0.089101 | 145.1359 | 963.2083 | 226.9527 | 0.0091981 | 0.015791 | 0.020548 | 1.0522 | 0.070009 | 0.0020211
963.8383 | 219.236 | 0.009425 | 0.015411 | 0.018726 | 1.0034 | 0.074275 | 0.003793 | 1572.6865 | 0.08919 | 145.1525 | 963.8383 | 226.871 | 0.0097532 | 0.015948 | 0.019378 | 1.0383 | 0.076862 | 0.0039251
964.4682 | 218.1516 | 0.009768 | 0.014607 | 0.018292 | 0.98031 | 0.065851 | -0.001855 | 1563.7894 | 0.08921 | 145.1672 | 964.4682 | 227.004 | 0.010164 | 0.0152 | 0.019034 | 1.0201 | 0.068523 | -0.0019303
965.0982 | 217.0143 | 0.009496 | 0.014516 | 0.018779 | 0.95468 | 0.056863 | -2.1e-05 | 1555.1545 | 0.089172 | 145.1804 | 965.0982 | 227.1935 | 0.0099414 | 0.015197 | 0.01966 | 0.99946 | 0.05953 | -2.1985e-05
965.7282 | 216.2566 | 0.008359 | 0.014555 | 0.016171 | 0.93807 | 0.0643 | 0.00702 | 1549.8598 | 0.089243 | 145.1928 | 965.7282 | 227.0119 | 0.0087747 | 0.015279 | 0.016975 | 0.98472 | 0.067498 | 0.0073691
966.3582 | 215.1883 | 0.008445 | 0.014791 | 0.016587 | 0.91736 | 0.068906 | -0.003626 | 1542.003 | 0.089245 | 145.205 | 966.3582 | 227.0561 | 0.0089107 | 0.015607 | 0.017502 | 0.96795 | 0.072706 | -0.003826
966.9882 | 213.8905 | 0.009196 | 0.014877 | 0.017994 | 0.89753 | 0.065582 | 0.004999 | 1533.5075 | 0.089273 | 145.2175 | 966.9882 | 226.8846 | 0.0097547 | 0.015781 | 0.019087 | 0.95206 | 0.069566 | 0.0053027
967.6182 | 212.6896 | 0.006988 | 0.014458 | 0.017358 | 0.88152 | 0.049217 | 0.00116 | 1524.7653 | 0.089238 | 145.231 | 967.6182 | 227.0151 | 0.0074587 | 0.015432 | 0.018527 | 0.94089 | 0.052532 | 0.0012381
968.2482 | 211.7807 | 0.006361 | 0.014932 | 0.014746 | 0.8608 | 0.061751 | 0.000493 | 1517.8283 | 0.089267 | 145.2459 | 968.2482 | 227.0275 | 0.0068189 | 0.016007 | 0.015808 | 0.92277 | 0.066197 | 0.00052849
968.8782 | 210.9261 | 0.010682 | 0.013309 | 0.016873 | 0.84044 | 0.064075 | 0.001405 | 1511.3371 | 0.089371 | 145.2629 | 968.8782 | 226.8431 | 0.011488 | 0.014313 | 0.018146 | 0.90386 | 0.06891 | 0.001511
969.5082 | 209.97 | 0.008525 | 0.013737 | 0.01604 | 0.82516 | 0.06471 | 0.0014 | 1504.6839 | 0.089305 | 145.2825 | 969.5082 | 227.0136 | 0.009217 | 0.014852 | 0.017342 | 0.89214 | 0.069963 | 0.0015136
970.1381 | 208.8673 | 0.007388 | 0.014538 | 0.014811 | 0.80448 | 0.070457 | 0.002991 | 1497.5804 | 0.089339 | 145.3053 | 970.1381 | 226.8403 | 0.0080237 | 0.015789 | 0.016085 | 0.8737 | 0.07652 | 0.0032484
970.7681 | 207.8823 | 0.008191 | 0.013635 | 0.014304 | 0.78996 | 0.059749 | 0.005945 | 1490.6732 | 0.089437 | 145.3316 | 970.7681 | 226.6106 | 0.0089289 | 0.014863 | 0.015593 | 0.86113 | 0.065132 | 0.0064806
971.3981 | 207.1203 | 0.006696 | 0.013672 | 0.013533 | 0.77 | 0.067179 | 0.005501 | 1486.077 | 0.089411 | 145.3612 | 971.3981 | 226.5892 | 0.0073254 | 0.014957 | 0.014805 | 0.84238 | 0.073494 | 0.0060181
972.0281 | 206.2447 | 0.006984 | 0.013785 | 0.014275 | 0.75519 | 0.063851 | 0.002573 | 1480.4953 | 0.089457 | 145.3936 | 972.0281 | 226.4155 | 0.007667 | 0.015133 | 0.015671 | 0.82905 | 0.070096 | 0.0028246
972.6581 | 205.4711 | 0.007936 | 0.013345 | 0.014881 | 0.74048 | 0.043507 | -0.000129 | 1474.604 | 0.089484 | 145.4285 | 972.6581 | 226.4525 | 0.0087464 | 0.014708 | 0.016401 | 0.81609 | 0.04795 | -0.00014217
973.2881 | 204.6067 | 0.006693 | 0.013 | 0.01101 | 0.72039 | 0.056039 | -0.001485 | 1469.5842 | 0.089454 | 145.4656 | 973.2881 | 226.4057 | 0.0074061 | 0.014385 | 0.012183 | 0.79714 | 0.062009 | -0.0016432
973.9181 | 203.874 | 0.009431 | 0.012567 | 0.013064 | 0.70763 | 0.054968 | 0.003747 | 1464.1448 | 0.089447 | 145.5044 | 973.9181 | 226.5103 | 0.010478 | 0.013962 | 0.014515 | 0.7862 | 0.061071 | 0.004163
974.5481 | 202.9753 | 0.008086 | 0.014085 | 0.012878 | 0.68673 | 0.061259 | 0.005348 | 1457.3768 | 0.089512 | 145.5447 | 974.5481 | 226.4574 | 0.0090215 | 0.015714 | 0.014368 | 0.76617 | 0.068346 | 0.0059667
975.178 | 201.9595 | 0.008069 | 0.01337 | 0.011571 | 0.67574 | 0.044887 | 0.004459 | 1450.5763 | 0.089484 | 145.5861 | 975.178 | 226.5157 | 0.0090501 | 0.014996 | 0.012978 | 0.75791 | 0.050345 | 0.0050012
975.808 | 201.1286 | 0.008153 | 0.013094 | 0.010592 | 0.6658 | 0.041543 | 0.005617 | 1445.5818 | 0.089492 | 145.6282 | 975.808 | 226.4078 | 0.0091777 | 0.01474 | 0.011923 | 0.74948 | 0.046764 | 0.006323
976.438 | 200.2923 | 0.007316 | 0.012612 | 0.014022 | 0.64912 | 0.045014 | -0.001848 | 1440.6198 | 0.089674 | 145.6706 | 976.438 | 225.8507 | 0.0082496 | 0.014221 | 0.015811 | 0.73195 | 0.050758 | -0.0020838
977.0681 | 199.8012 | 0.00801 | 0.013055 | 0.013369 | 0.6413 | 0.064351 | 0.004063 | 1436.7723 | 0.089647 | 145.713 | 977.0681 | 226.0333 | 0.0090616 | 0.014769 | 0.015124 | 0.7255 | 0.0728 | 0.0045964
977.698 | 199.3405 | 0.007939 | 0.013227 | 0.012584 | 0.62454 | 0.04514 | -0.000466 | 1434.0636 | 0.089659 | 145.7551 | 977.698 | 225.9744 | 0.0089997 | 0.014994 | 0.014265 | 0.70798 | 0.051171 | -0.00052826
978.328 | 198.495 | 0.006694 | 0.01259 | 0.010745 | 0.61006 | 0.048601 | -0.000468 | 1428.0355 | 0.089727 | 145.7964 | 978.328 | 225.8584 | 0.0076168 | 0.014326 | 0.012226 | 0.69416 | 0.055301 | -0.00053252
978.958 | 197.8242 | 0.005959 | 0.012894 | 0.012723 | 0.59484 | 0.058857 | 0.005877 | 1423.211 | 0.089683 | 145.8367 | 978.958 | 226.0307 | 0.0068087 | 0.014732 | 0.014537 | 0.67965 | 0.067249 | 0.006715
979.588 | 196.8922 | 0.006194 | 0.012885 | 0.011736 | 0.58409 | 0.039077 | 0.005195 | 1416.9534 | 0.08973 | 145.8755 | 979.588 | 225.9007 | 0.0071066 | 0.014783 | 0.013465 | 0.67015 | 0.044834 | 0.0059604
980.2136 | 195.8836 | 0.005969 | 0.012803 | 0.010419 | 0.5751 | 0.043661 | -0.004539 | 1410.8285 | 0.089662 | 145.9123 | 980.2136 | 225.9487 | 0.0068851 | 0.014768 | 0.012018 | 0.66337 | 0.050362 | -0.0052357
980.8392 | 195.0791 | 0.00712 | 0.013593 | 0.011934 | 0.56132 | 0.048831 | 0.002734 | 1405.5856 | 0.089774 | 145.9471 | 980.8392 | 225.6307 | 0.0082351 | 0.015722 | 0.013803 | 0.64923 | 0.056478 | 0.0031622
981.4648 | 194.2529 | 0.006891 | 0.011769 | 0.011668 | 0.54717 | 0.039256 | 0.001852 | 1400.7542 | 0.089817 | 145.9799 | 981.4648 | 225.3924 | 0.0079957 | 0.013656 | 0.013538 | 0.63489 | 0.045549 | 0.0021489
982.0904 | 193.6403 | 0.006662 | 0.012205 | 0.010412 | 0.53977 | 0.043272 | 0.002329 | 1396.495 | 0.089896 | 146.0105 | 982.0904 | 225.2158 | 0.0077483 | 0.014195 | 0.01211 | 0.62779 | 0.050328 | 0.0027088
982.716 | 192.9454 | 0.006432 | 0.012271 | 0.00942 | 0.52983 | 0.043907 | 0.005739 | 1392.0483 | 0.089875 | 146.039 | 982.716 | 225.2216 | 0.007508 | 0.014324 | 0.010996 | 0.61846 | 0.051252 | 0.006699
983.3416 | 192.177 | 0.006813 | 0.011772 | 0.00928 | 0.5149 | 0.06041 | -0.001977 | 1386.6163 | 0.08985 | 146.0654 | 983.3416 | 225.3067 | 0.0079875 | 0.013801 | 0.01088 | 0.60367 | 0.070824 | -0.0023178
983.9672 | 191.3446 | 0.005562 | 0.012873 | 0.009726 | 0.51101 | 0.052529 | 0.000964 | 1381.6349 | 0.089912 | 146.0895 | 983.9672 | 225.0218 | 0.0065409 | 0.015139 | 0.011438 | 0.60094 | 0.061774 | 0.0011337
984.5928 | 190.1838 | 0.006452 | 0.012886 | 0.007935 | 0.49254 | 0.054274 | 0.001865 | 1373.9177 | 0.089907 | 146.1114 | 984.5928 | 224.9596 | 0.0076318 | 0.015242 | 0.0093859 | 0.58261 | 0.064198 | 0.002206
985.2184 | 189.0019 | 0.006269 | 0.011779 | 0.009627 | 0.48354 | 0.043566 | 0.006194 | 1366.9769 | 0.089924 | 146.1309 | 985.2184 | 224.6843 | 0.0074525 | 0.014003 | 0.011445 | 0.57483 | 0.051791 | 0.0073634
985.844 | 188.2625 | 0.00731 | 0.012705 | 0.010529 | 0.47666 | 0.039655 | 0.004173 | 1362.4659 | 0.090056 | 146.1481 | 985.844 | 224.2435 | 0.0087071 | 0.015133 | 0.012541 | 0.56775 | 0.047234 | 0.0049705
986.4696 | 187.3971 | 0.007892 | 0.012191 | 0.00906 | 0.46429 | 0.046508 | -0.003969 | 1355.9565 | 0.090098 | 146.163 | 986.4696 | 224.2012 | 0.009442 | 0.014585 | 0.010839 | 0.55548 | 0.055642 | -0.0047485
987.0952 | 186.9795 | 0.006844 | 0.012535 | 0.009239 | 0.45502 | 0.021625 | 0.006923 | 1351.7219 | 0.090001 | 146.1754 | 987.0952 | 224.6639 | 0.0082234 | 0.015061 | 0.011101 | 0.54672 | 0.025983 | 0.0083183
987.7208 | 186.5756 | 0.006558 | 0.012668 | 0.009812 | 0.4471 | 0.055124 | 0.000344 | 1347.1707 | 0.090134 | 146.1853 | 987.7208 | 224.6186 | 0.0078952 | 0.015251 | 0.011813 | 0.53827 | 0.066364 | 0.00041414
988.3464 | 185.9276 | 0.006017 | 0.012458 | 0.009395 | 0.437 | 0.046101 | 0.003285 | 1342.3525 | 0.090083 | 146.1927 | 988.3464 | 224.7805 | 0.0072744 | 0.015061 | 0.011358 | 0.52833 | 0.055735 | 0.0039715
988.972 | 185.0415 | 0.00522 | 0.011847 | 0.009044 | 0.43226 | 0.034809 | 0.000101 | 1337.4265 | 0.090143 | 146.1976 | 988.972 | 224.3926 | 0.0063301 | 0.014366 | 0.010967 | 0.52419 | 0.042212 | 0.00012248
989.5976 | 183.9978 | 0.005137 | 0.01227 | 0.00856 | 0.42317 | 0.048473 | -0.001493 | 1329.2205 | 0.090178 | 146.1999 | 989.5976 | 224.4209 | 0.0062656 | 0.014966 | 0.010441 | 0.51614 | 0.059122 | -0.001821
990.2232 | 182.9301 | 0.005564 | 0.012145 | 0.006692 | 0.40796 | 0.041156 | -0.002619 | 1323.5976 | 0.090143 | 146.1996 | 990.2232 | 224.1531 | 0.0068178 | 0.014882 | 0.0082 | 0.49989 | 0.05043 | -0.0032092
990.8488 | 182.2707 | 0.007419 | 0.012385 | 0.007592 | 0.40626 | 0.034392 | 0.000123 | 1319.1329 | 0.090206 | 146.1967 | 990.8488 | 223.9387 | 0.009115 | 0.015216 | 0.0093276 | 0.49913 | 0.042254 | 0.00015112
991.4744 | 181.7687 | 0.006214 | 0.011997 | 0.008358 | 0.39665 | 0.038396 | -0.001207 | 1316.6613 | 0.090311 | 146.1914 | 991.4744 | 223.473 | 0.0076397 | 0.01475 | 0.010276 | 0.48765 | 0.047205 | -0.0014839
992.1 | 181.287 | 0.005519 | 0.012248 | 0.005241 | 0.38985 | 0.044668 | 0.004699 | 1313.6282 | 0.090384 | 146.184 | 992.1 | 223.2053 | 0.0067951 | 0.01508 | 0.0064529 | 0.47999 | 0.054996 | 0.0057855
992.7256 | 180.8586 | 0.005179 | 0.012328 | 0.00726 | 0.38075 | 0.041323 | -0.005516 | 1311.0586 | 0.090352 | 146.1747 | 992.7256 | 223.1781 | 0.0063908 | 0.015213 | 0.0089588 | 0.46984 | 0.050992 | -0.0068067
993.3512 | 180.5733 | 0.005861 | 0.011517 | 0.007041 | 0.37585 | 0.041367 | 0.00014 | 1309.1694 | 0.090394 | 146.1637 | 993.3512 | 223.028 | 0.007239 | 0.014225 | 0.0086964 | 0.46422 | 0.051093 | 0.00017292
993.9768 | 179.8064 | 0.007 | 0.01185 | 0.008272 | 0.36738 | 0.036867 | -0.004401 | 1303.2736 | 0.090401 | 146.1513 | 993.9768 | 223.0487 | 0.0086835 | 0.0147 | 0.010261 | 0.45573 | 0.045733 | -0.0054594
994.6024 | 178.6959 | 0.003652 | 0.017034 | 0.008051 | 0.35545 | 0.035219 | -0.004648 | 1295.7733 | 0.090452 | 146.1376 | 994.6024 | 222.8077 | 0.0045535 | 0.021239 | 0.010038 | 0.4432 | 0.043913 | -0.0057954
995.228 | 177.5454 | 0.00556 | 0.013696 | 0.007107 | 0.34876 | 0.032974 | 0.002152 | 1289.4389 | 0.090387 | 146.1228 | 995.228 | 222.5967 | 0.0069708 | 0.017171 | 0.0089104 | 0.43726 | 0.041341 | 0.0026981
995.8536 | 176.6724 | 0.005579 | 0.011808 | 0.007215 | 0.34703 | 0.025092 | -0.001935 | 1284.2214 | 0.090455 | 146.1073 | 995.8536 | 222.2133 | 0.0070171 | 0.014852 | 0.0090748 | 0.43649 | 0.03156 | -0.0024338
996.4792 | 176.0932 | 0.005496 | 0.013035 | 0.007716 | 0.33599 | 0.045539 | -0.000575 | 1280.4104 | 0.090533 | 146.0912 | 996.4792 | 221.928 | 0.0069265 | 0.016428 | 0.0097244 | 0.42344 | 0.057392 | -0.00072467
997.1048 | 175.5656 | 0.005159 | 0.013305 | 0.007362 | 0.33081 | 0.037643 | 0.001262 | 1277.2336 | 0.090612 | 146.0747 | 997.1048 | 221.5933 | 0.0065115 | 0.016793 | 0.0092921 | 0.41753 | 0.047512 | 0.0015929
997.7304 | 175.1289 | 0.005229 | 0.012743 | 0.009972 | 0.32728 | 0.033713 | -0.005285 | 1273.8747 | 0.09057 | 146.058 | 997.7304 | 221.7042 | 0.0066196 | 0.016132 | 0.012624 | 0.41432 | 0.042679 | -0.0066905
998.356 | 174.7092 | 0.005859 | 0.011757 | 0.009622 | 0.3198 | 0.027509 | -0.000283 | 1272.0523 | 0.090654 | 146.0415 | 998.356 | 221.2591 | 0.0074201 | 0.01489 | 0.012186 | 0.40501 | 0.034839 | -0.0003584
998.9816 | 174.0912 | 0.005521 | 0.011503 | 0.008018 | 0.31856 | 0.029253 | -0.003898 | 1267.6013 | 0.090611 | 146.0252 | 998.9816 | 221.3315 | 0.0070191 | 0.014624 | 0.010194 | 0.40501 | 0.037191 | -0.0049557
999.6072 | 172.7759 | 0.003758 | 0.012049 | 0.008723 | 0.30789 | 0.043456 | -0.003882 | 1258.486 | 0.090652 | 146.0095 | 999.6072 | 221.1256 | 0.0048096 | 0.015421 | 0.011164 | 0.39406 | 0.055617 | -0.0049683
1000.2328 | 171.557 | 0.005155 | 0.012065 | 0.008108 | 0.29939 | 0.028193 | 0.002714 | 1249.8253 | 0.090664 | 145.9945 | 1000.2328 | 221.0346 | 0.0066417 | 0.015545 | 0.010446 | 0.38573 | 0.036324 | 0.0034967
1000.8584 | 170.6337 | 0.003539 | 0.012016 | 0.008547 | 0.28413 | 0.025408 | 0.00726 | 1243.6977 | 0.090842 | 145.9806 | 1000.8584 | 220.4737 | 0.0045727 | 0.015526 | 0.011043 | 0.36712 | 0.032829 | 0.0093806
1001.484 | 169.8702 | 0.005751 | 0.011433 | 0.006026 | 0.28532 | 0.02941 | 0.001593 | 1239.7717 | 0.090767 | 145.9679 | 1001.484 | 220.3465 | 0.0074599 | 0.01483 | 0.0078166 | 0.3701 | 0.038149 | 0.0020664
1002.1096 | 169.4913 | 0.006384 | 0.011838 | 0.006203 | 0.28281 | 0.024348 | 0.006574 | 1237.1051 | 0.090808 | 145.957 | 1002.1096 | 220.2126 | 0.0082944 | 0.015381 | 0.0080593 | 0.36745 | 0.031634 | 0.0085413
1002.7352 | 169.1317 | 0.00528 | 0.011939 | 0.006117 | 0.27661 | 0.049357 | -0.003626 | 1234.9281 | 0.090806 | 145.948 | 1002.7352 | 220.125 | 0.0068719 | 0.015539 | 0.0079613 | 0.36 | 0.064238 | -0.0047192
1003.3608 | 168.5928 | 0.004941 | 0.011504 | 0.005636 | 0.27053 | 0.029001 | -0.00655 | 1230.6033 | 0.090853 | 145.9414 | 1003.3608 | 220.0702 | 0.0064497 | 0.015017 | 0.0073569 | 0.35313 | 0.037856 | -0.0085499
1003.9864 | 167.9893 | 0.006437 | 0.011637 | 0.006533 | 0.26419 | 0.03923 | 0.00368 | 1226.9411 | 0.09087 | 145.9374 | 1003.9864 | 219.8896 | 0.0084257 | 0.015232 | 0.0085514 | 0.34581 | 0.05135 | 0.0048169
1004.612 | 166.8413 | 0.004618 | 0.011783 | 0.005913 | 0.25734 | 0.032463 | 0.003703 | 1219.2511 | 0.090883 | 145.9364 | 1004.612 | 219.7312 | 0.0060819 | 0.015518 | 0.0077875 | 0.33892 | 0.042754 | 0.0048769
1005.2376 | 165.7004 | 0.004634 | 0.010785 | 0.007666 | 0.25291 | 0.02344 | -0.002413 | 1211.0798 | 0.090942 | 145.9387 | 1005.2376 | 219.5617 | 0.0061403 | 0.014291 | 0.010158 | 0.33512 | 0.031059 | -0.0031974
1005.8632 | 164.599 | 0.005621 | 0.010266 | 0.005468 | 0.24402 | 0.038218 | -0.005139 | 1204.2509 | 0.090904 | 145.9447 | 1005.8632 | 219.4411 | 0.0074938 | 0.013686 | 0.0072899 | 0.32532 | 0.050952 | -0.0068512
1006.4888 | 164.0156 | 0.008138 | 0.010529 | 0.007417 | 0.2424 | 0.032598 | 0.000982 | 1201.4599 | 0.090967 | 145.9547 | 1006.4888 | 219.0334 | 0.010868 | 0.014061 | 0.009905 | 0.32371 | 0.043533 | 0.0013114
1007.1144 | 163.5022 | 0.004279 | 0.012017 | 0.005486 | 0.23274 | 0.040586 | 0.002123 | 1198.1116 | 0.091054 | 145.9689 | 1007.1144 | 218.7697 | 0.0057254 | 0.016079 | 0.0073404 | 0.31141 | 0.054305 | 0.0028406
1007.74 | 163.2122 | 0.002663 | 0.011451 | 0.005327 | 0.23392 | 0.04632 | 0.005758 | 1196.1968 | 0.091152 | 145.9879 | 1007.74 | 218.5257 | 0.0035655 | 0.015332 | 0.0071324 | 0.3132 | 0.062018 | 0.0077094
1008.3656 | 162.8647 | 0.00431 | 0.011358 | 0.007664 | 0.22797 | 0.029353 | 0.002581 | 1194.6588 | 0.09105 | 146.0117 | 1008.3656 | 218.6217 | 0.0057855 | 0.015246 | 0.010288 | 0.30601 | 0.039402 | 0.0034646
1008.9912 | 162.2241 | 0.004479 | 0.012285 | 0.006512 | 0.22557 | 0.035636 | -0.004001 | 1190.4838 | 0.091106 | 146.0409 | 1008.9912 | 218.4343 | 0.006031 | 0.016542 | 0.0087684 | 0.30373 | 0.047984 | -0.0053873
1009.6168 | 161.4402 | 0.004545 | 0.011749 | 0.007009 | 0.21782 | 0.055512 | -0.001748 | 1184.5781 | 0.091157 | 146.0758 | 1009.6168 | 218.3911 | 0.0061483 | 0.015894 | 0.0094815 | 0.29466 | 0.075095 | -0.0023646
1010.2424 | 160.2141 | 0.004253 | 0.011559 | 0.007108 | 0.21721 | 0.036828 | 0.003229 | 1176.8209 | 0.09119 | 146.1165 | 1010.2424 | 218.1443 | 0.0057908 | 0.015739 | 0.0096781 | 0.29574 | 0.050144 | 0.0043965
1010.868 | 159.4538 | 0.003705 | 0.010731 | 0.004574 | 0.21251 | 0.028355 | 0.006425 | 1170.7575 | 0.091234 | 146.1626 | 1010.868 | 218.1966 | 0.0050699 | 0.014684 | 0.0062591 | 0.2908 | 0.038801 | 0.008792
1011.4936 | 158.6037 | 0.004485 | 0.010956 | 0.003554 | 0.2054 | 0.032377 | 0.010749 | 1165.3583 | 0.091346 | 146.212 | 1011.4936 | 217.8437 | 0.0061602 | 0.015048 | 0.0048815 | 0.28212 | 0.04447 | 0.014764
1012.1192 | 158.1518 | 0.003171 | 0.011201 | 0.006809 | 0.20071 | 0.039217 | 0.005307 | 1163.0201 | 0.091214 | 146.263 | 1012.1192 | 218.0524 | 0.004372 | 0.015443 | 0.0093879 | 0.27673 | 0.054071 | 0.007317
1012.7448 | 157.7484 | 0.005533 | 0.011826 | 0.006179 | 0.19895 | 0.029633 | 0.004181 | 1160.688 | 0.091214 | 146.3133 | 1012.7448 | 218.0067 | 0.0076466 | 0.016343 | 0.0085393 | 0.27494 | 0.040953 | 0.0057781
1013.3704 | 157.3497 | 0.004528 | 0.012122 | 0.005291 | 0.19693 | 0.039309 | -0.002398 | 1157.9869 | 0.091313 | 146.3611 | 1013.3704 | 217.8 | 0.0062676 | 0.016779 | 0.0073237 | 0.27258 | 0.054411 | -0.0033193
1013.996 | 156.8163 | 0.004746 | 0.011575 | 0.00894 | 0.19081 | 0.047857 | -0.004889 | 1154.6728 | 0.091403 | 146.4044 | 1013.996 | 217.5341 | 0.0065836 | 0.016057 | 0.012401 | 0.2647 | 0.066387 | -0.006782
1014.6216 | 156.0639 | 0.005017 | 0.011398 | 0.006204 | 0.19312 | 0.047329 | -0.00463 | 1149.1303 | 0.091399 | 146.4412 | 1014.6216 | 217.5989 | 0.0069952 | 0.015892 | 0.0086502 | 0.26927 | 0.065991 | -0.0064556
1015.2472 | 155.0027 | 0.002837 | 0.011423 | 0.002609 | 0.18113 | 0.026933 | -0.0012 | 1141.4466 | 0.091489 | 146.4695 | 1015.2472 | 217.401 | 0.0039791 | 0.016021 | 0.0036593 | 0.25405 | 0.037775 | -0.0016831
1015.8728 | 153.9038 | 0.005045 | 0.011345 | 0.002559 | 0.18074 | 0.034915 | 0.002454 | 1135.1834 | 0.091516 | 146.4873 | 1015.8728 | 217.0123 | 0.0071137 | 0.015997 | 0.0036083 | 0.25486 | 0.049232 | 0.0034603
1016.4984 | 153.1451 | 0.004141 | 0.010996 | 0.005472 | 0.1764 | 0.024738 | 0.004968 | 1129.8967 | 0.091515 | 146.4927 | 1016.4984 | 216.9635 | 0.0058666 | 0.015578 | 0.0077523 | 0.24991 | 0.035047 | 0.0070383
1017.124 | 152.5709 | 0.004003 | 0.010512 | 0.003709 | 0.17141 | 0.029284 | -0.0048 | 1126.4909 | 0.091467 | 146.4836 | 1017.124 | 216.9058 | 0.005691 | 0.014945 | 0.005273 | 0.24368 | 0.041632 | -0.006824
1017.7496 | 152.3633 | 0.003553 | 0.011528 | 0.005964 | 0.17202 | 0.023087 | -0.001197 | 1125.0924 | 0.091555 | 146.4581 | 1017.7496 | 216.6315 | 0.0050517 | 0.016391 | 0.0084797 | 0.24459 | 0.032825 | -0.0017019
1018.3752 | 152.0577 | 0.005247 | 0.010684 | 0.006571 | 0.17124 | 0.044661 | -0.001457 | 1123.1121 | 0.091548 | 146.4142 | 1018.3752 | 216.531 | 0.0074718 | 0.015214 | 0.0093571 | 0.24384 | 0.063598 | -0.0020748
1019.0008 | 151.2129 | 0.002403 | 0.011394 | 0.003232 | 0.16369 | 0.029955 | -0.001263 | 1116.5638 | 0.091604 | 146.3499 | 1019.0008 | 216.3626 | 0.0034383 | 0.016303 | 0.0046245 | 0.23422 | 0.042861 | -0.0018072
1019.6264 | 150.3432 | 0.004761 | 0.011177 | 0.003839 | 0.16278 | 0.034548 | 0.007566 | 1110.344 | 0.091421 | 146.2632 | 1019.6264 | 216.6284 | 0.0068601 | 0.016105 | 0.0055316 | 0.23454 | 0.04978 | 0.010902
1020.252 | 149.3894 | 0.005279 | 0.011031 | 0.006023 | 0.15931 | 0.034028 | 0.004847 | 1104.2362 | 0.091509 | 146.1523 | 1020.252 | 216.0725 | 0.0076354 | 0.015955 | 0.0087115 | 0.23042 | 0.049217 | 0.0070106
1020.8776 | 148.6657 | 0.00534 | 0.011558 | 0.004656 | 0.15432 | 0.034067 | 0.002835 | 1097.9764 | 0.091677 | 146.0175 | 1020.8776 | 215.6559 | 0.0077463 | 0.016766 | 0.006754 | 0.22386 | 0.049418 | 0.0041125
1021.5032 | 148.113 | 0.004432 | 0.011244 | 0.004664 | 0.15251 | 0.034102 | 0.005146 | 1094.3106 | 0.091699 | 145.8625 | 1021.5032 | 215.2942 | 0.0064423 | 0.016344 | 0.0067795 | 0.22169 | 0.04957 | 0.0074801
1022.1288 | 147.7695 | 0.003219 | 0.012624 | 0.003622 | 0.15044 | 0.019955 | 0.001539 | 1091.9764 | 0.091694 | 145.691 | 1022.1288 | 215.0129 | 0.0046838 | 0.018369 | 0.0052702 | 0.2189 | 0.029036 | 0.0022393
1022.7544 | 147.3504 | 0.004503 | 0.010762 | 0.00238 | 0.1508 | 0.032466 | 0.004033 | 1089.9752 | 0.09167 | 145.5069 | 1022.7544 | 214.5817 | 0.0065576 | 0.015672 | 0.0034659 | 0.21961 | 0.047279 | 0.0058731
1023.38 | 146.9578 | 0.005427 | 0.011484 | 0.003847 | 0.1481 | 0.023989 | 0.008541 | 1087.5617 | 0.091699 | 145.3139 | 1023.38 | 214.1322 | 0.0079077 | 0.016733 | 0.0056055 | 0.2158 | 0.034954 | 0.012445
1024.0056 | 146.2435 | 0.00569 | 0.011639 | 0.003591 | 0.14247 | 0.025745 | -0.001244 | 1083.0477 | 0.091787 | 145.1158 | 1024.0056 | 213.4816 | 0.0083061 | 0.01699 | 0.005242 | 0.20797 | 0.037582 | -0.001816
1024.6312 | 145.3869 | 0.002178 | 0.011723 | 0.005976 | 0.14042 | 0.030874 | 0.002626 | 1078.0248 | 0.091878 | 144.9164 | 1024.6312 | 212.7167 | 0.0031866 | 0.017152 | 0.0087435 | 0.20544 | 0.045172 | 0.0038421
1025.2568 | 144.3695 | 0.003359 | 0.011967 | 0.004607 | 0.13797 | 0.023539 | -7.4e-05 | 1071.4656 | 0.091795 | 144.7195 | 1025.2568 | 212.4249 | 0.0049424 | 0.017608 | 0.0067787 | 0.20301 | 0.034635 | -0.00010888
1025.8824 | 143.4955 | 0.005049 | 0.012333 | 0.006397 | 0.13579 | 0.025284 | -0.001403 | 1065.342 | 0.091865 | 144.5288 | 1025.8824 | 211.9113 | 0.0074563 | 0.018213 | 0.009447 | 0.20053 | 0.037339 | -0.0020719
1026.5081 | 142.7793 | 0.003582 | 0.011608 | 0.006082 | 0.13068 | 0.04287 | -0.001155 | 1060.5808 | 0.091844 | 144.3481 | 1026.5081 | 211.584 | 0.0053081 | 0.017202 | 0.0090129 | 0.19366 | 0.063529 | -0.0017116
1027.1335 | 142.1389 | 0.002674 | 0.010756 | 0.004323 | 0.13068 | 0.030443 | -0.005477 | 1054.9898 | 0.091908 | 144.1813 | 1027.1335 | 211.3598 | 0.0039762 | 0.015994 | 0.0064283 | 0.19431 | 0.045269 | -0.0081443
1027.7592 | 141.4199 | 0.0036 | 0.011141 | 0.005858 | 0.13169 | 0.030476 | -0.003938 | 1049.306 | 0.091828 | 144.0321 | 1027.7592 | 211.3934 | 0.0053813 | 0.016653 | 0.0087565 | 0.19686 | 0.045555 | -0.0058865
1028.3848 | 140.8071 | 0.004019 | 0.012133 | 0.004825 | 0.12583 | 0.03221 | 0.001234 | 1045.7576 | 0.091899 | 143.9043 | 1028.3848 | 210.8418 | 0.006018 | 0.018168 | 0.0072249 | 0.18841 | 0.048231 | 0.0018478
1029.0104 | 140.0339 | 0.00454 | 0.01246 | 0.005573 | 0.12187 | 0.025447 | -0.001953 | 1040.1391 | 0.091895 | 143.8017 | 1029.0104 | 210.6746 | 0.0068302 | 0.018746 | 0.0083843 | 0.18335 | 0.038284 | -0.0029382
1029.636 | 139.337 | 0.002305 | 0.012183 | 0.002362 | 0.11716 | 0.022077 | 0.004414 | 1035.4553 | 0.09188 | 143.728 | 1029.636 | 210.5009 | 0.0034822 | 0.018405 | 0.0035683 | 0.17699 | 0.033352 | 0.0066684
1030.2616 | 138.618 | 0.002006 | 0.011225 | 0.00765 | 0.11855 | 0.019274 | 0.003315 | 1030.2112 | 0.092111 | 143.687 | 1030.2616 | 209.8945 | 0.0030375 | 0.016997 | 0.011584 | 0.17951 | 0.029185 | 0.0050196
1030.8872 | 137.6685 | 0.002879 | 0.011884 | 0.003785 | 0.11458 | 0.02895 | -0.004387 | 1024.0332 | 0.091951 | 143.6793 | 1030.8872 | 210.0671 | 0.004393 | 0.018134 | 0.0057755 | 0.17484 | 0.044175 | -0.0066941
1031.5128 | 136.955 | 0.003245 | 0.011874 | 0.005714 | 0.11304 | 0.016511 | 0.001296 | 1019.237 | 0.092021 | 143.7016 | 1031.5128 | 209.8357 | 0.0049718 | 0.018193 | 0.0087547 | 0.1732 | 0.025297 | 0.0019857
1032.1384 | 136.563 | 0.00193 | 0.011241 | 0.005332 | 0.10665 | 0.017113 | -0.006408 | 1016.6229 | 0.092109 | 143.7505 | 1032.1384 | 209.6438 | 0.0029628 | 0.017257 | 0.0081854 | 0.16372 | 0.026271 | -0.0098372
1032.764 | 136.4785 | 0.002396 | 0.011602 | 0.003962 | 0.10829 | 0.033581 | -0.005949 | 1017.5021 | 0.092103 | 143.8225 | 1032.764 | 209.4499 | 0.0036771 | 0.017805 | 0.0060804 | 0.16619 | 0.051536 | -0.0091298
1033.3896 | 136.4593 | 0.004391 | 0.012468 | 0.003315 | 0.10343 | 0.009787 | -0.001618 | 1018.0393 | 0.091949 | 143.9142 | 1033.3896 | 209.7944 | 0.0067508 | 0.019168 | 0.0050965 | 0.15901 | 0.015047 | -0.0024875
1034.0153 | 135.9405 | 0.004189 | 0.01216 | 0.005164 | 0.10073 | 0.027343 | 0.002485 | 1014.1782 | 0.091997 | 144.0221 | 1034.0153 | 209.8408 | 0.0064662 | 0.01877 | 0.0079713 | 0.15548 | 0.042207 | 0.0038359
1034.6407 | 135.3676 | 0.002922 | 0.012422 | 0.007149 | 0.10045 | 0.021709 | -0.003391 | 1010.4039 | 0.092091 | 144.1428 | 1034.6407 | 209.6994 | 0.0045265 | 0.019243 | 0.011075 | 0.15561 | 0.03363 | -0.005253
1035.2664 | 134.5618 | 0.004303 | 0.012602 | 0.005247 | 0.099031 | 0.021734 | -0.001337 | 1004.8906 | 0.09205 | 144.2727 | 1035.2664 | 209.8762 | 0.0067114 | 0.019655 | 0.0081838 | 0.15446 | 0.033899 | -0.0020853
1035.892 | 133.8753 | 0.003593 | 0.011919 | 0.007039 | 0.095581 | 0.024054 | 0.004794 | 999.8418 | 0.092204 | 144.4085 | 1035.892 | 209.7058 | 0.0056282 | 0.01867 | 0.011026 | 0.14972 | 0.037679 | 0.0075094
1036.5176 | 133.3915 | 0.002681 | 0.012923 | 0.002373 | 0.099146 | 0.012188 | 0.000498 | 996.7112 | 0.092186 | 144.5467 | 1036.5176 | 209.8469 | 0.0042177 | 0.02033 | 0.0037331 | 0.15597 | 0.019174 | 0.00078344
1037.1432 | 133.0915 | 0.001875 | 0.0129 | 0.004433 | 0.092379 | 0.020746 | 0.001617 | 995.1936 | 0.092259 | 144.6838 | 1037.1432 | 209.7274 | 0.0029547 | 0.020328 | 0.0069856 | 0.14557 | 0.032692 | 0.0025481
1037.7688 | 132.7958 | 0.002187 | 0.013252 | 0.004706 | 0.088931 | 0.019673 | -0.000418 | 993.4778 | 0.092136 | 144.8164 | 1037.7688 | 210.0951 | 0.00346 | 0.020966 | 0.0074453 | 0.1407 | 0.031124 | -0.00066131
1038.3944 | 132.3879 | 0.004283 | 0.012811 | 0.005379 | 0.088291 | 0.017458 | 0.003899 | 992.5377 | 0.092231 | 144.941 | 1038.3944 | 209.6122 | 0.0067814 | 0.020284 | 0.0085167 | 0.13979 | 0.027642 | 0.0061734
1039.02 | 131.6194 | 0.003373 | 0.012692 | 0.004074 | 0.084594 | 0.022597 | 0.008432 | 987.1724 | 0.092177 | 145.0541 | 1039.02 | 209.8136 | 0.0053769 | 0.020232 | 0.0064943 | 0.13485 | 0.036022 | 0.013441
1039.6456 | 130.7169 | 0.002459 | 0.01227 | 0.003227 | 0.086504 | 0.014707 | 0.005485 | 980.047 | 0.092178 | 145.1523 | 1039.6456 | 210.0293 | 0.003951 | 0.019715 | 0.005185 | 0.13899 | 0.02363 | 0.008813
1040.2712 | 130.1058 | 0.00313 | 0.01361 | 0.004557 | 0.083823 | 0.009652 | 0.003889 | 973.2535 | 0.092276 | 145.2322 | 1040.2712 | 210.3999 | 0.0050617 | 0.022009 | 0.0073693 | 0.13555 | 0.015609 | 0.0062891
1040.8969 | 129.7257 | 0.001349 | 0.012335 | 0.002525 | 0.080256 | 0.006866 | 0.000489 | 968.6266 | 0.092338 | 145.2926 | 1040.8969 | 210.7322 | 0.0021914 | 0.020038 | 0.0041017 | 0.13037 | 0.011153 | 0.00079435
1041.5225 | 129.5771 | 0.003708 | 0.012823 | 0.003063 | 0.081785 | 0.023346 | -0.001558 | 967.8085 | 0.092303 | 145.3349 | 1041.5225 | 210.8105 | 0.0060326 | 0.020862 | 0.0049832 | 0.13306 | 0.037982 | -0.0025347
1042.1479 | 129.5775 | 0.001828 | 0.012053 | 0.003729 | 0.08255 | 0.018303 | -0.001089 | 968.8314 | 0.092396 | 145.3605 | 1042.1479 | 210.4132 | 0.0029684 | 0.019572 | 0.0060553 | 0.13405 | 0.029721 | -0.0017684
1042.7736 | 128.95 | 0.002394 | 0.013013 | 0.005448 | 0.083068 | 0.032545 | 0.003242 | 966.1548 | 0.092406 | 145.3711 | 1042.7736 | 209.9674 | 0.0038981 | 0.021189 | 0.0088709 | 0.13526 | 0.052993 | 0.0052789
1043.3992 | 128.1405 | 0.003468 | 0.013164 | 0.001961 | 0.079509 | 0.019517 | 0.003933 | 961.2553 | 0.092429 | 145.3681 | 1043.3992 | 209.6558 | 0.0056741 | 0.021538 | 0.0032085 | 0.13009 | 0.031933 | 0.0064349
1044.0248 | 127.3562 | 0.003574 | 0.012597 | 0.005395 | 0.074922 | 0.025779 | -0.007609 | 956.3912 | 0.09248 | 145.3531 | 1044.0248 | 209.297 | 0.0058735 | 0.020702 | 0.0088661 | 0.12313 | 0.042365 | -0.012505
1044.6504 | 126.8909 | 0.000979 | 0.013255 | 0.003957 | 0.075956 | 0.025791 | -0.001695 | 953.1376 | 0.092358 | 145.3275 | 1044.6504 | 209.4837 | 0.0016162 | 0.021883 | 0.0065326 | 0.1254 | 0.042578 | -0.0027983
1045.276 | 126.3165 | 0.001496 | 0.013568 | 0.003442 | 0.074443 | 0.02807 | -0.003727 | 949.1161 | 0.092598 | 145.2931 | 1045.276 | 208.8265 | 0.0024732 | 0.022431 | 0.0056903 | 0.12307 | 0.046405 | -0.0061615
1045.9016 | 125.6568 | 0.003133 | 0.013594 | 0.003981 | 0.074969 | 0.011079 | -0.001237 | 945.4203 | 0.092474 | 145.2511 | 1045.9016 | 208.7663 | 0.0052052 | 0.022585 | 0.006614 | 0.12455 | 0.018407 | -0.0020552
1046.5272 | 125.0285 | 0.000794 | 0.012763 | 0.003011 | 0.067082 | 0.015618 | -0.001249 | 940.583 | 0.092568 | 145.2033 | 1046.5272 | 208.5101 | 0.0013242 | 0.021285 | 0.0050214 | 0.11187 | 0.026046 | -0.002083
1047.1528 | 124.4372 | 0.002487 | 0.013992 | 0.003821 | 0.068628 | 0.028106 | -0.004204 | 935.3826 | 0.092657 | 145.151 | 1047.1528 | 208.4014 | 0.0041651 | 0.023433 | 0.0063992 | 0.11493 | 0.047071 | -0.0070407
1047.7784 | 123.3259 | 0.002292 | 0.013653 | 0.001602 | 0.067623 | 0.009421 | 0.000331 | 927.2344 | 0.092649 | 145.0959 | 1047.7784 | 208.2955 | 0.0038712 | 0.02306 | 0.0027058 | 0.11421 | 0.015912 | 0.00055905
1048.4041 | 122.0156 | 0.005621 | 0.013793 | 0.003733 | 0.062285 | 0.019649 | -0.002612 | 917.8031 | 0.09279 | 145.0395 | 1048.4041 | 207.8022 | 0.009573 | 0.023491 | 0.0063576 | 0.10608 | 0.033464 | -0.0044484
1049.0295 | 121.1285 | 0.002518 | 0.0132 | 0.003691 | 0.06166 | 0.023613 | -0.002608 | 912.2388 | 0.092877 | 144.9833 | 1049.0295 | 207.2743 | 0.0043088 | 0.022588 | 0.006316 | 0.10551 | 0.040406 | -0.0044628
1049.6552 | 120.5644 | 0.001509 | 0.013318 | 0.001542 | 0.062816 | 0.028151 | -0.000569 | 908.579 | 0.09293 | 144.9288 | 1049.6552 | 206.9441 | 0.0025901 | 0.02286 | 0.0026468 | 0.10782 | 0.04832 | -0.00097667