Draft test plan for MOBY NET Stability Testing at NIST

The Stability Source modifications and NIST Characterization of Monitor System due June 2016. Project ends July 2017; both dates are from original proposal.

Major Tasks.

| Tasks/Date | 2015 |  |  |  |  |  |  |  |  |  | 2016 |  |  |  |  |  |  |  | 2017 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 1 | 2 | 3 | 4 | 5 | 6 | Q3 | Q4 | Q1 | Q2 | Q3 |
| RSL Preparations |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Laboratory modifications | x | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Stability and repeatability of RSL assets |  |  | x | x | x | x | x | x | x | x | x | x | x | x | x | x |  |  |  |  |  |
| Develop & produce DAQ and DA software (1) |  | x | x | x | x | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Implement LabView DAQ for Temp/RH monitor |  |  | x | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Order CAS fiber optic; finalize foreoptic | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| SQM delivered to UM |  | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| SQM Mounting plate for foreoptic and fiducials |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Design |  | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Produce |  |  | x | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Deliver to NIST |  |  |  |  | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| System testing at UM (2) |  | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Ship SQM and CAS to NIST |  |  |  | x |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Radiometric Characterization at NIST |  |  |  |  | x | x | x | x | x | x | x | x | x | x | x | x |  |  |  |  |  |
| Radiometric Calibration at NIST |  |  |  |  |  |  |  | x | x |  |  |  |  |  | x | x |  | x |  | x |  |
| Shipping tests |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | x | x | x | x | x |

- DAQ of CAS – in conjunction with John Woodward, Tom Larason at NIST; MLML team; DA (data analysis) software should have an eye to the final product in MOBY-Net – means Stephanie’s MLDbase variables. Send sample NIST files for CAS, simulated SQM files, and sample Temp/RH files to everyone NOW.

- Ken, if you have a reproducible and stable source and radiometer you could take measurements with the CAS and of the SQM and these would serve as an initial baseline for stability with shipping. We’d need to repeat at UM with the same sources/radiometer at some future date.

CAS140CT-156 Operation – Data Acquisition (DAQ) & Data Analysis (DA)

Table 1. CAS parameters, settings, and controls (for Vaux, Faux). A Vaux parameter is measured after each shutter open or shutter closed scan. A Faux parameter is measured once before and once after the dark, light, dark radiometric data set. However, I think one can put Faux data into the Vaux? Help, Steph. The list is in no particular order. Uses USB interface.

| Parameter | Nominal or Range | Available from CAS via software? | Vaux or Faux |
| --- | --- | --- | --- |
| Serial number & name, M/N | alpha/numeric | Yes | F |
| CAS_DLL version number | character 64 | Yes | F |
| Config and Calibration files | string | Yes | F |
| S/N of config and calibration files | should match S/N of unit | Yes (flag) | F |
| Is Dark Current Auto sampled? |  | Yes (flag) | V |
| Are we using Dark Current array? |  | Set flag | V |
| Amp offset for last dark current | numeric | Yes | V |
| Is the LED switched off for meas? |  | Set flag | V |
| Pixel to wavelength conversion | (pxl) | Yes | F |
| Dead or visible pixels | nnn | Yes | F |
| set a saturation value in DN | 15 bit total | Flag | F |
| CCD temperature(1) | -10C (resolution is 0.5C) | Yes | V |
| Integration time | 10 msec to 60 sec | Yes | V |
| Set shutter(2) | N/A | Yes | N/A |
| Get shutter position | -1, 0, 1 or -1010 | Yes | V |
| Range filter wheel | open or OD 1 to OD 4 | Yes | V |
| Order sorting filters | are they on the CCD? | ? | ? |
| Accumulation(3) | nnn | Yes | V |
| CIE color space parameters | varies | Yes | V |
| Collect Mode(4) | 0, 1, 2 | No | V |
| If Mode=0, what type of dark?(5) | 0, 1, …n | Yes & No | V |
| Fiber adaptor | PLG-320 | No | F |
| Optical fiber | OFG-xxx (for example) | No | F |
| Foreoptic | MOBY Lu (for example) | No | F |
| Distance to source | xx cm | No | F |
| Laser excitation wavelength(6) | xxx.xx nm | No | V |
| Source observed | SQM-5002 (for example) | No | F |
| Source Level | LO Bank (for example) | No | F |
| Operator | Johnson (for example) | No | F |
| Location name | NIST RSL (for example) | No | F |
| Location latitude | 39.135 | No | F |
| Location longitude | -77.217 | No | F |
| Date& time stamp | YYMMDD.HHMMSS | No | V |
| Calibration experiment? | Y or N | No | F |
| Characterization experiment? | Y or N | No | F |
| Routine operation? | Y or N | No | F |
| Comment field | character limit? | No | F |
| Name of this data file | decide format | No | F |
| Type of this data file(7) | nnn | No | F |
| Ambient temp & RH(8) | 35C max | No | V |

- turn off CAS if this goes below -8C – requires Operator Intervention – build Notice into Labview DAQ

- We have control of the shutter in software but the existing NIST programs don’t save the dark scans

- defined as number of scans per CCD read

- refers to darks(0) with a lens cap or source shuttered or off, ambients(1) with an on axis opaque disk, or signal(2) from unobstructed source;

- if mode=0 for darks, are these lens cap, source shuttered, or CAS darks (camera shuttered)?

- only applies for laser characterization experiments

- a numeric code identifies this as the output of a radiometer, the MOBY-NET CAS in particular, when measuring a source

- ambient temperature and humidity can come from various sources, from hand notes to a logging Fluke on the GPIB network; turn CAS off if temp is above 35C – requires Operator Intervention – build Notice into Labview DAQ if possible

We are assuming an acquisition sequence as in MOBY: Close shutter, take darks; open shutter, take lights, close shutter, take darks – choice of M dark scans and N light scans. I am confused about all the options with Dark Current in the CAS, and which command can get us DN (typically reports DN/sec). There are dead pixels and also different types of commands for reading data (lights).

SQM-5002 – Data Acquisition (DAQ) & Data Analysis (DA)

Table 2. SQM-5002 parameters, settings, and controls (for Vaux, Faux). A Vaux parameter is measured repeatedly. A Faux parameter is measured once for each “file”. The list is in no particular order. Uses RS232 interface.

| Parameter | Nominal or Range | Available from SQM via software? | Vaux or Faux |
| --- | --- | --- | --- |
| Serial number & name, M/N | alpha/numeric | No | F |
| SQM Manager version | alpha/numeric | No | F |
| Lamp set description(1) | alpha/numeric | No | F |
| LO lamp set current | numeric | Yes | V |
| LO lamp current(2) | numeric | Yes | V |
| LO register (factory diagnostic) | numeric | Yes | V |
| HI lamp set current | numeric | Yes | V |
| HI lamp current | numeric | Yes | V |
| HI register (factory diagnostic) | numeric | Yes | V |
| Red monitor | numeric | Yes | V |
| Blue monitor | numeric | Yes | V |
| White monitor(3) | numeric | Yes | V |
| Case temperature (near lamp rings) | numeric | Yes | V |
| Monitor detector mounting block temperature | numeric | Yes | V |
| LO & HI lamp shunt resistors temperature(4) | numeric | Yes | V |
| Date, time observer connected AC – starts STANDBY mode(5) | dd-mon-yyyy 24:xx | No | F |
| Date, time observer switched STANDBY to PREHEAT(6) | dd-mon-yyyy 24:xx | No | F |
| Date, time observer switched Lamps On (any set) | dd-mon-yyyy 24:xx | No | F |
| Date, time observer switched Lamps Off (any set) | dd-mon-yyyy 24:xx | No | F |
| Date, time observer switched from LAMP ON to PREHEAT | dd-mon-yyyy 24:xx | No | F |
| Date, time observer switched from LAMP ON to STANDBY | dd-mon-yyyy 24:xx | No | F |
| Date, time observer unplugged SQM from AC | dd-mon-yyyy 24:xx | No | F |
| LO lamp hours cumulative (or this file only?)(7) | HH.hh | No | F |
| HI lamp hours cumulative (or this file only?) | HH.hh | No | F |
| Adaptor flange configuration(8) | 0, 1, …n | No | F |
| Radiometer(s) observing | MOBY-NET CAS (for example) | No | F |
| Source Level | LO Bank (for example) | No | F |
| Operator | Johnson (for example) | No | F |
| Location name | NIST RSL (for example) | No | F |
| Location latitude | 39.135 | No | F |
| Location longitude | -77.217 | No | F |
| Date& time stamp(see 5) | YYMMDD.HHMMSS | No | F |
| Calibration experiment? | Y or N | No | F |
| Characterization experiment? | Y or N | No | F |
| Routine operation? | Y or N | No | F |
| Comment field | character limit? | No | F |
| Name of this data file(9) | decide format | No | F |
| Type of this data file(10) | nnn | No | F |
| Ambient temp & RH(11) | 33C max | No | F |

- keep track of the LO and HI bank of lamps, indicate which lamp was replaced and when

- we will have to analyze the data to figure out of LO, HI, or BOTH are one; also record in log sheet

- can we add a filter (UV or photopic) to the unfiltered (white) monitor detector?

- the shunt resistor is held at 45C, it is heated only and cannot be cooled

- there are no date or times in the SQM_Monitor log file, and no header information

- this has to be on a log sheet or logsheet GUI

- if cumulative, we’ll need to load a cfg files with hours to date?

- that is, what is on the SQM exit aperture – a shutter closed/open, a MOBY-NET adaptor, nothing, etc…

- the file name is typed in SQM-Manager, but it should be automated

- a numeric code identifies this as the output of a radiometer, the MOBY-NET CAS in particular, when measuring a source

- ambient temperature and humidity can come from various sources, from hand notes to a logging Fluke on the GPIB network; turn SQM off if temp is above 33C, RH approached 100%, or direct contact with precipitation of spray is possible.

Strawman for Final Report – for DAQ and DA software design

here I am going to suggest some figures and tables for the final report, and this should guide the experiment sequence.